## spec_table
id | sku | spec_value | spec_name
1 | 1097901 | 600 | Maximum Energy Density (millijoules per square centimeter)
2 | 1097901 | 250 | Minimum Wavelength (nanometers)
3 | 1097901 | 11000 | Maximum Wavelength (nanometers)
4 | 1097901 | 0.01 | Minimum Power (watts)
5 | 1097901 | 10 | Maximum Power (watts)
6 | 1097901 | 10 | Maximum Long-Pulse Energy (joules)
7 | 1097901 | Air | Cooling Type
8 | 1097901 | 0.25 - 11 | Wavelength Range (micrometers)
9 | 1097901 | 10 mW to 10W | Power Range
10 | 1097901 | 30 | Maximum Intermittent Power (watts)(<5 minutes)
11 | 1097901 | 0.5 to 10 | Long-Pulse Joules Range (joules)
12 | 1097901 | 1 | Resolution (milliwatts)
13 | 1097901 | 6 | Maximum Power Density (kilowatts per square centimeter)
14 | 1097901 | Broadband | Detector Coating
15 | 1097901 | ±1 | Calibration Uncertainty (%)(k=2)
16 | 1097901 | ±1 | Power Linearity (%)
17 | 1097901 | ±1.5 | Spectral Compensation Accuracy (%)
18 | 1097901 | 514 | Calibration Wavelength (nanometers)
19 | 1097901 | Air | Cooling Method
20 | 1097901 | 2 | Cable Length (meters)
21 | 1097901 | None | Window
22 | 1097901 | None | Diffuser
23 | 1097901 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
24 | 1097901 | PM DB-25 | Cable Type
25 | 1097901 | 19 mm | Active Area Diameter
26 | 1097901 | 0.2 | Noise Equivalent Power (milliwatts)
27 | 1097901 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
28 | 1098293 | CE, RoHS, WEEE, ISO 17025 | Compliance
30 | 1098293 | Optical | Sensor Compatibility
31 | 1098293 | Power | Measurement Type
32 | 1098293 | No | Thermopile sensor compatibility
33 | 1098293 | Yes | Optical sensor compatibility
34 | 1098293 | No | Pyroelectric sensor compatibility
35 | 1098293 | No | PowerMax-Pro sensor compatibility
36 | 1098293 | No | USB Interface
37 | 1098293 | No | RS232 Interface
38 | 1098293 | No | Analog output
39 | 1098293 | No | External Trigger Input
40 | 1098293 | Fixed Segment LCD | Display Type
41 | 1098293 | Yes | Battery Operation
42 | 1098293 | 10 uW to 1W | Measurement Range
43 | 1098293 | 0.01 uW | Measurement Resolution
44 | 1098293 | N/A | Maximum Pulse Energy Repetition Rate (hertz)
45 | 1098293 | 10 Hz | Power Sampling Rate
46 | 1098293 | ±5.0% of reading | Accuracy - System
47 | 1098293 | N/A | Accuracy - Analog Output
48 | 1098293 | ±8.0% | Calibration Uncertainty (%)(k=2)
49 | 1098293 | 3-digit fixed-segment LCD | Display
50 | 1098293 | None | Statistical Analysis
51 | 1098293 | None | Computer Interface
52 | 1098293 | N/A | Analog Output Text
53 | 1098293 | N/A | Trigger
54 | 1098293 | Battery operated | Instrument Power
55 | 1098293 | 168 x 23 x 30 mm (6.6 x 0.9 x 0.7 in.) | Size (H x W x D)
56 | 1098293 | 44 g (0.09 lbs.) | Weight
57 | 1098293 | Yes | C24
58 | 1098297 | 26 x 89 mm (1.02 x 3.5 in.), fixed-segment LCD and analog tuning needle | Display
59 | 1098297 | 12VDC - center positive; (90 to 260 VAC, 50/60 Hz AC power adapter included) | Instrument Power
60 | 1098297 | CE, RoHS, WEEE, ISO 17025 | Compliance
62 | 1098297 | Optical | Sensor Compatibility
63 | 1098297 | Thermopile (PM-model) | Sensor Compatibility
64 | 1098297 | Power | Measurement Type
65 | 1098297 | Analog output | Outputs & Interfaces
66 | 1098297 | Yes | Thermopile sensor compatibility
67 | 1098297 | Yes | Optical sensor compatibility
68 | 1098297 | No | Pyroelectric sensor compatibility
69 | 1098297 | No | PowerMax-Pro sensor compatibility
70 | 1098297 | No | USB Interface
71 | 1098297 | No | RS232 Interface
72 | 1098297 | Yes | Analog output
73 | 1098297 | No | External Trigger Input
74 | 1098297 | Fixed Segment LCD | Display Type
75 | 1098297 | Yes | Battery Operation
76 | 1098297 | Sensor dependent (reference sensor specifications) | Measurement Range
77 | 1098297 | 0.1% of full-scale for 10s scale; 0.3% of full-scale range for 3s scale | Measurement Resolution
78 | 1098297 | N/A | Maximum Pulse Energy Repetition Rate (hertz)
79 | 1098297 | 20 Hz | Power Sampling Rate
80 | 1098297 | Meter accuracy + sensor accuracy | Accuracy - System
81 | 1098297 | ±1.0% | Accuracy - Analog Output
82 | 1098297 | ±1.0% | Calibration Uncertainty (%)(k=2)
83 | 1098297 | None | Statistical Analysis
84 | 1098297 | None | Computer Interface
85 | 1098297 | 0 to 2 VDC | Analog Output Text
86 | 1098297 | N/A | Trigger
87 | 1098297 | 193 x 117 x 46 mm (7.6 x 4.6 x 1.8 in.) | Size (H x W x D)
88 | 1098297 | 0.8 kg (1.8 lbs.) | Weight
89 | 1098297 | Yes | C24
90 | 1098314 | 600 | Maximum Energy Density (millijoules per square centimeter)
91 | 1098314 | 250 | Minimum Wavelength (nanometers)
92 | 1098314 | 11000 | Maximum Wavelength (nanometers)
93 | 1098314 | 0.1 | Minimum Power (watts)
94 | 1098314 | 30 | Maximum Power (watts)
95 | 1098314 | 50 | Maximum Long-Pulse Energy (joules)
96 | 1098314 | Air | Cooling Type
97 | 1098314 | 0.25 - 11 | Wavelength Range (micrometers)
98 | 1098314 | 100 mW to 30W | Power Range
99 | 1098314 | 50 | Maximum Intermittent Power (watts)(<5 minutes)
100 | 1098314 | 0.5 to 50 | Long-Pulse Joules Range (joules)
101 | 1098314 | 10 | Resolution (milliwatts)
102 | 1098314 | 6 | Maximum Power Density (kilowatts per square centimeter)
103 | 1098314 | Broadband | Detector Coating
104 | 1098314 | ±1 | Calibration Uncertainty (%)(k=2)
105 | 1098314 | ±1 | Power Linearity (%)
106 | 1098314 | ±1.5 | Spectral Compensation Accuracy (%)
107 | 1098314 | 514 | Calibration Wavelength (nanometers)
108 | 1098314 | Air | Cooling Method
109 | 1098314 | 2 | Cable Length (meters)
110 | 1098314 | None | Window
111 | 1098314 | None | Diffuser
112 | 1098314 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
113 | 1098314 | PM DB-25 | Cable Type
114 | 1098314 | 19 mm | Active Area Diameter
115 | 1098314 | 0.5 | Noise Equivalent Power (milliwatts)
116 | 1098314 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
117 | 1098329 | 600 | Maximum Energy Density (millijoules per square centimeter)
118 | 1098329 | 250 | Minimum Wavelength (nanometers)
119 | 1098329 | 11000 | Maximum Wavelength (nanometers)
120 | 1098329 | 0.01 | Minimum Power (watts)
121 | 1098329 | 2 | Maximum Power (watts)
122 | 1098329 | 2 | Maximum Long-Pulse Energy (joules)
123 | 1098329 | Air | Cooling Type
124 | 1098329 | 0.25 - 11 | Wavelength Range (micrometers)
125 | 1098329 | 10 mW to 2W | Power Range
126 | 1098329 | 5 | Maximum Intermittent Power (watts)(<5 minutes)
127 | 1098329 | 0.5 to 2 | Long-Pulse Joules Range (joules)
128 | 1098329 | 1 | Resolution (milliwatts)
129 | 1098329 | 6 | Maximum Power Density (kilowatts per square centimeter)
130 | 1098329 | Broadband | Detector Coating
131 | 1098329 | ±1 | Calibration Uncertainty (%)(k=2)
132 | 1098329 | ±1 | Power Linearity (%)
133 | 1098329 | ±1.5 | Spectral Compensation Accuracy (%)
134 | 1098329 | 514 | Calibration Wavelength (nanometers)
135 | 1098329 | Air | Cooling Method
136 | 1098329 | 2 | Cable Length (meters)
137 | 1098329 | None | Window
138 | 1098329 | None | Diffuser
139 | 1098329 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
140 | 1098329 | PM DB-25 | Cable Type
141 | 1098329 | 19 mm | Active Area Diameter
142 | 1098329 | 0.2 | Noise Equivalent Power (milliwatts)
143 | 1098329 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
144 | 1098336 | 300 | Minimum Wavelength (nanometers)
145 | 1098336 | 11000 | Maximum Wavelength (nanometers)
146 | 1098336 | 0.0005 | Minimum Power (watts)
147 | 1098336 | 2 | Maximum Power (watts)
148 | 1098336 | Air | Cooling Type
149 | 1098336 | 0.19 to 11 | Wavelength Range (micrometers)
150 | 1098336 | 500 µW to 2W | Power Range
151 | 1098336 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
152 | 1098336 | 50 | Maximum Energy Density (millijoules per square centimeter)
153 | 1098336 | 0.05 | Resolution (milliwatts)
154 | 1098336 | 0.5 | Maximum Power Density (kilowatts per square centimeter)
155 | 1098336 | Black | Detector Coating
156 | 1098336 | ±1 | Calibration Uncertainty (%)(k=2)
157 | 1098336 | ±1 | Power Linearity (%)
158 | 1098336 | ±1.5 | Spectral Compensation Accuracy (%)
159 | 1098336 | 514 | Calibration Wavelength (nanometers)
160 | 1098336 | Air | Cooling Method
161 | 1098336 | 2 | Cable Length (meters)
162 | 1098336 | None | Window
163 | 1098336 | None | Diffuser
164 | 1098336 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
165 | 1098336 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
166 | 1098336 | PM DB-25 | Cable Type
167 | 1098336 | 19 mm | Active Area Diameter
168 | 1098336 | 0.02 | Noise Equivalent Power (milliwatts)
169 | 1098336 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
170 | 1098341 | 300 | Minimum Wavelength (nanometers)
171 | 1098341 | 2000 | Maximum Wavelength (nanometers)
172 | 1098341 | 0.0001 | Minimum Power (watts)
173 | 1098341 | 1 | Maximum Power (watts)
174 | 1098341 | 1 | Maximum Long-Pulse Energy (joules)
175 | 1098341 | Air | Cooling Type
176 | 1098341 | 0.3 - 2 | Wavelength Range (micrometers)
177 | 1098341 | 100 µW to 1W | Power Range
178 | 1098341 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
179 | 1098341 | .001 to 1 | Long-Pulse Joules Range (joules)
180 | 1098341 | 0.01 | Resolution (milliwatts)
181 | 1098341 | 0.5 | Maximum Power Density (kilowatts per square centimeter)
182 | 1098341 | Black | Detector Coating
183 | 1098341 | ±1 | Calibration Uncertainty (%)(k=2)
184 | 1098341 | ±1 | Power Linearity (%)
185 | 1098341 | ±1.5 | Spectral Compensation Accuracy (%)
186 | 1098341 | 514 | Calibration Wavelength (nanometers)
187 | 1098341 | Air | Cooling Method
188 | 1098341 | 2 | Cable Length (meters)
189 | 1098341 | Yes | Window
190 | 1098341 | None | Diffuser
191 | 1098341 | 50 | Maximum Energy Density (millijoules per square centimeter)
192 | 1098341 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
193 | 1098341 | PM DB-25 | Cable Type
194 | 1098341 | 19 mm | Active Area Diameter
195 | 1098341 | 0.003 | Noise Equivalent Power (milliwatts)
196 | 1098341 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
197 | 1098350 | 300 | Minimum Wavelength (nanometers)
198 | 1098350 | 11000 | Maximum Wavelength (nanometers)
199 | 1098350 | 0.0001 | Minimum Power (watts)
200 | 1098350 | 1 | Maximum Power (watts)
201 | 1098350 | 1 | Maximum Long-Pulse Energy (joules)
202 | 1098350 | Air | Cooling Type
203 | 1098350 | 0.19 - 11 | Wavelength Range (micrometers)
204 | 1098350 | 100 µW to 1W | Power Range
205 | 1098350 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
206 | 1098350 | .001 to 1 | Long-Pulse Joules Range (joules)
207 | 1098350 | 0.01 | Resolution (milliwatts)
208 | 1098350 | 0.5 | Maximum Power Density (kilowatts per square centimeter)
209 | 1098350 | Black | Detector Coating
210 | 1098350 | ±1 | Calibration Uncertainty (%)(k=2)
211 | 1098350 | ±1 | Power Linearity (%)
212 | 1098350 | ±1.5 | Spectral Compensation Accuracy (%)
213 | 1098350 | 514 | Calibration Wavelength (nanometers)
214 | 1098350 | Air | Cooling Method
215 | 1098350 | 2 | Cable Length (meters)
216 | 1098350 | None | Window
217 | 1098350 | None | Diffuser
218 | 1098350 | 50 | Maximum Energy Density (millijoules per square centimeter)
219 | 1098350 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
220 | 1098350 | PM DB-25 | Cable Type
221 | 1098350 | 10 mm | Active Area Diameter
222 | 1098350 | 0.003 | Noise Equivalent Power (milliwatts)
223 | 1098350 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
224 | 1098398 | 600 | Maximum Energy Density (millijoules per square centimeter)
225 | 1098398 | 250 | Minimum Wavelength (nanometers)
226 | 1098398 | 11000 | Maximum Wavelength (nanometers)
227 | 1098398 | 0.3 | Minimum Power (watts)
228 | 1098398 | 150 | Maximum Power (watts)
229 | 1098398 | 150 | Maximum Long-Pulse Energy (joules)
230 | 1098398 | Air | Cooling Type
231 | 1098398 | 0.25 - 11 | Wavelength Range (micrometers)
232 | 1098398 | 300 mW to 150W | Power Range
233 | 1098398 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
234 | 1098398 | 1 to 150 | Long-Pulse Joules Range (joules)
235 | 1098398 | 30 | Resolution (milliwatts)
236 | 1098398 | 6 | Maximum Power Density (kilowatts per square centimeter)
237 | 1098398 | Broadband | Detector Coating
238 | 1098398 | ±1 | Calibration Uncertainty (%)(k=2)
239 | 1098398 | ±1 | Power Linearity (%)
240 | 1098398 | ±1.5 | Spectral Compensation Accuracy (%)
241 | 1098398 | 514 | Calibration Wavelength (nanometers)
242 | 1098398 | Air | Cooling Method
243 | 1098398 | 2 | Cable Length (meters)
244 | 1098398 | None | Window
245 | 1098398 | None | Diffuser
246 | 1098398 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
247 | 1098398 | PM DB-25 | Cable Type
248 | 1098398 | 50 mm | Active Area Diameter
249 | 1098398 | 1 | Noise Equivalent Power (milliwatts)
250 | 1098398 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
251 | 1098400 | 300 | Minimum Wavelength (nanometers)
252 | 1098400 | 2000 | Maximum Wavelength (nanometers)
253 | 1098400 | 0.0001 | Minimum Power (watts)
254 | 1098400 | 1 | Maximum Power (watts)
255 | 1098400 | 1 | Maximum Long-Pulse Energy (joules)
256 | 1098400 | Air | Cooling Type
257 | 1098400 | 0.3 - 2 | Wavelength Range (micrometers)
258 | 1098400 | 100 µW to 1W | Power Range
259 | 1098400 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
260 | 1098400 | .001 to 1 | Long-Pulse Joules Range (joules)
261 | 1098400 | 0.01 | Resolution (milliwatts)
262 | 1098400 | 0.5 | Maximum Power Density (kilowatts per square centimeter)
263 | 1098400 | Black | Detector Coating
264 | 1098400 | ±1 | Calibration Uncertainty (%)(k=2)
265 | 1098400 | ±1 | Power Linearity (%)
266 | 1098400 | ±1.5 | Spectral Compensation Accuracy (%)
267 | 1098400 | 514 | Calibration Wavelength (nanometers)
268 | 1098400 | Air | Cooling Method
269 | 1098400 | 2 | Cable Length (meters)
270 | 1098400 | Yes | Window
271 | 1098400 | None | Diffuser
272 | 1098400 | 50 | Maximum Energy Density (millijoules per square centimeter)
273 | 1098400 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
274 | 1098400 | PM DB-25 | Cable Type
275 | 1098400 | 10 mm | Active Area Diameter
276 | 1098400 | 0.003 | Noise Equivalent Power (milliwatts)
277 | 1098400 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
278 | 1098407 | 600 | Maximum Energy Density (millijoules per square centimeter)
279 | 1098407 | 250 | Minimum Wavelength (nanometers)
280 | 1098407 | 11000 | Maximum Wavelength (nanometers)
281 | 1098407 | 0.3 | Minimum Power (watts)
282 | 1098407 | 150 | Maximum Power (watts)
283 | 1098407 | 150 | Maximum Long-Pulse Energy (joules)
284 | 1098407 | Air | Cooling Type
285 | 1098407 | 0.25 - 11 | Wavelength Range (micrometers)
286 | 1098407 | 300 mW to 150W | Power Range
287 | 1098407 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
288 | 1098407 | 1 to 150 | Long-Pulse Joules Range (joules)
289 | 1098407 | 30 | Resolution (milliwatts)
290 | 1098407 | 6 | Maximum Power Density (kilowatts per square centimeter)
291 | 1098407 | Broadband | Detector Coating
292 | 1098407 | ±1 | Calibration Uncertainty (%)(k=2)
293 | 1098407 | ±1 | Power Linearity (%)
294 | 1098407 | ±1.5 | Spectral Compensation Accuracy (%)
295 | 1098407 | 514 | Calibration Wavelength (nanometers)
296 | 1098407 | Air | Cooling Method
297 | 1098407 | 2 | Cable Length (meters)
298 | 1098407 | None | Window
299 | 1098407 | None | Diffuser
300 | 1098407 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
301 | 1098407 | PM DB-25 | Cable Type
302 | 1098407 | 19 mm | Active Area Diameter
303 | 1098407 | 1 | Noise Equivalent Power (milliwatts)
304 | 1098407 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
305 | 1098412 | 600 | Maximum Energy Density (millijoules per square centimeter)
306 | 1098412 | 250 | Minimum Wavelength (nanometers)
307 | 1098412 | 11000 | Maximum Wavelength (nanometers)
308 | 1098412 | 0.3 | Minimum Power (watts)
309 | 1098412 | 150 | Maximum Power (watts)
310 | 1098412 | 150 | Maximum Long-Pulse Energy (joules)
311 | 1098412 | 0.25 - 11 | Wavelength Range (micrometers)
312 | 1098412 | 300 mW to 150W | Power Range
313 | 1098412 | 80 | Maximum Intermittent Power (watts)(<5 minutes)
314 | 1098412 | 1 to 150 | Long-Pulse Joules Range (joules)
315 | 1098412 | 30 | Resolution (milliwatts)
316 | 1098412 | 6 | Maximum Power Density (kilowatts per square centimeter)
317 | 1098412 | Broadband | Detector Coating
318 | 1098412 | ±1 | Calibration Uncertainty (%)(k=2)
319 | 1098412 | ±1 | Power Linearity (%)
320 | 1098412 | ±1.5 | Spectral Compensation Accuracy (%)
321 | 1098412 | 514 | Calibration Wavelength (nanometers)
322 | 1098412 | Water | Cooling Method
323 | 1098412 | 2 | Cable Length (meters)
324 | 1098412 | None | Window
325 | 1098412 | None | Diffuser
326 | 1098412 | Water, Air (intermittent) | Cooling Type
327 | 1098412 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
328 | 1098412 | PM DB-25 | Cable Type
329 | 1098412 | 50 mm | Active Area Diameter
330 | 1098412 | 1 | Noise Equivalent Power (milliwatts)
331 | 1098412 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
332 | 1098413 | 300 | Minimum Wavelength (nanometers)
333 | 1098413 | 11000 | Maximum Wavelength (nanometers)
334 | 1098413 | 0.0001 | Minimum Power (watts)
335 | 1098413 | 1 | Maximum Power (watts)
336 | 1098413 | 1 | Maximum Long-Pulse Energy (joules)
337 | 1098413 | Air | Cooling Type
338 | 1098413 | 0.3 - 11 | Wavelength Range (micrometers)
339 | 1098413 | 100 µW to 1W | Power Range
340 | 1098413 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
341 | 1098413 | .001 to 1 | Long-Pulse Joules Range (joules)
342 | 1098413 | 0.01 | Resolution (milliwatts)
343 | 1098413 | 0.5 | Maximum Power Density (kilowatts per square centimeter)
344 | 1098413 | Black | Detector Coating
345 | 1098413 | ±1 | Calibration Uncertainty (%)(k=2)
346 | 1098413 | ±1 | Power Linearity (%)
347 | 1098413 | ±1.5 | Spectral Compensation Accuracy (%)
348 | 1098413 | 514 | Calibration Wavelength (nanometers)
349 | 1098413 | Air | Cooling Method
350 | 1098413 | 2 | Cable Length (meters)
351 | 1098413 | None | Window
352 | 1098413 | None | Diffuser
353 | 1098413 | 50 | Maximum Energy Density (millijoules per square centimeter)
354 | 1098413 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
355 | 1098413 | PM DB-25 | Cable Type
356 | 1098413 | 19 mm | Active Area Diameter
357 | 1098413 | 0.005 | Noise Equivalent Power (milliwatts)
358 | 1098413 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
359 | 1098417 | 600 | Maximum Energy Density (millijoules per square centimeter)
360 | 1098417 | 250 | Minimum Wavelength (nanometers)
361 | 1098417 | 11000 | Maximum Wavelength (nanometers)
362 | 1098417 | 1 | Minimum Power (watts)
363 | 1098417 | 300 | Maximum Power (watts)
364 | 1098417 | 300 | Maximum Long-Pulse Energy (joules)
365 | 1098417 | Fan | Cooling Type
366 | 1098417 | 0.25 to 11 | Wavelength Range (micrometers)
367 | 1098417 | 1W to 300W | Power Range
368 | 1098417 | 450 | Maximum Intermittent Power (watts)(<5 minutes)
369 | 1098417 | 1 to 300 | Long-Pulse Joules Range (joules)
370 | 1098417 | 100 | Resolution (milliwatts)
371 | 1098417 | 6 | Maximum Power Density (kilowatts per square centimeter)
372 | 1098417 | Broadband | Detector Coating
373 | 1098417 | ±1 | Calibration Uncertainty (%)(k=2)
374 | 1098417 | ±1 | Power Linearity (%)
375 | 1098417 | ±1.5 | Spectral Compensation Accuracy (%)
376 | 1098417 | 514 | Calibration Wavelength (nanometers)
377 | 1098417 | Fan | Cooling Method
378 | 1098417 | 2 | Cable Length (meters)
379 | 1098417 | None | Window
380 | 1098417 | None | Diffuser
381 | 1098417 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
382 | 1098417 | PM DB-25 | Cable Type
383 | 1098417 | 50 mm | Active Area Diameter
384 | 1098417 | 5 mW (w/o fan), 100 mW (with fan) | Noise Equivalent Power (milliwatts)
385 | 1098417 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
386 | 1098443 | 80, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
387 | 1098443 | 600 | Maximum Energy Density (millijoules per square centimeter)
388 | 1098443 | Water, Air (intermittent) | Cooling Type
389 | 1098443 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
390 | 1098443 | 150 | Minimum Wavelength (nanometers)
391 | 1098443 | 1000 | Maximum Wavelength (nanometers)
392 | 1098443 | 0.3 | Minimum Power (watts)
393 | 1098443 | 150 | Maximum Power (watts)
394 | 1098443 | 150 | Maximum Long-Pulse Energy (joules)
395 | 1098443 | 0.15 - 1 | Wavelength Range (micrometers)
396 | 1098443 | 300 mW to 150W | Power Range
397 | 1098443 | 1 to 150 | Long-Pulse Joules Range (joules)
398 | 1098443 | 30 | Resolution (milliwatts)
399 | 1098443 | 6 | Maximum Power Density (kilowatts per square centimeter)
400 | 1098443 | ±1 | Calibration Uncertainty (%)(k=2)
401 | 1098443 | ±1 | Power Linearity (%)
402 | 1098443 | ±1.5 | Spectral Compensation Accuracy (%)
403 | 1098443 | 514 | Calibration Wavelength (nanometers)
404 | 1098443 | Water | Cooling Method
405 | 1098443 | 2 | Cable Length (meters)
406 | 1098443 | None | Window
407 | 1098443 | None | Diffuser
408 | 1098443 | PM DB-25 | Cable Type
409 | 1098443 | 50 mm | Active Area Diameter
410 | 1098443 | 1 | Noise Equivalent Power (milliwatts)
411 | 1098443 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
412 | 1098443 | UV | Detector Coating
413 | 1098455 | 600 | Maximum Energy Density (millijoules per square centimeter)
414 | 1098455 | 150 | Minimum Wavelength (nanometers)
415 | 1098455 | 1000 | Maximum Wavelength (nanometers)
416 | 1098455 | 0.3 | Minimum Power (watts)
417 | 1098455 | 150 | Maximum Power (watts)
418 | 1098455 | 150 | Maximum Long-Pulse Energy (joules)
419 | 1098455 | Air | Cooling Type
420 | 1098455 | 0.15 - 1 | Wavelength Range (micrometers)
421 | 1098455 | 300 mW to 150W | Power Range
422 | 1098455 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
423 | 1098455 | 1 to 150 | Long-Pulse Joules Range (joules)
424 | 1098455 | 30 | Resolution (milliwatts)
425 | 1098455 | 6 | Maximum Power Density (kilowatts per square centimeter)
426 | 1098455 | UV | Detector Coating
427 | 1098455 | ±1 | Calibration Uncertainty (%)(k=2)
428 | 1098455 | ±1 | Power Linearity (%)
429 | 1098455 | ±1.5 | Spectral Compensation Accuracy (%)
430 | 1098455 | 514 | Calibration Wavelength (nanometers)
431 | 1098455 | Air | Cooling Method
432 | 1098455 | 2 | Cable Length (meters)
433 | 1098455 | None | Window
434 | 1098455 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
435 | 1098455 | None | Diffuser
436 | 1098455 | PM DB-25 | Cable Type
437 | 1098455 | 50 mm | Active Area Diameter
438 | 1098455 | 1 | Noise Equivalent Power (milliwatts)
439 | 1098455 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
440 | 1098457 | 600 | Maximum Energy Density (millijoules per square centimeter)
441 | 1098457 | 150 | Minimum Wavelength (nanometers)
442 | 1098457 | 1000 | Maximum Wavelength (nanometers)
443 | 1098457 | 0.01 | Minimum Power (watts)
444 | 1098457 | 2 | Maximum Power (watts)
445 | 1098457 | 2 | Maximum Long-Pulse Energy (joules)
446 | 1098457 | Air | Cooling Type
447 | 1098457 | 0.15 - 1 | Wavelength Range (micrometers)
448 | 1098457 | 10 mW to 2W | Power Range
449 | 1098457 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
450 | 1098457 | 0.5 to 2 | Long-Pulse Joules Range (joules)
451 | 1098457 | 1 | Resolution (milliwatts)
452 | 1098457 | 6 | Maximum Power Density (kilowatts per square centimeter)
453 | 1098457 | UV | Detector Coating
454 | 1098457 | ±1 | Calibration Uncertainty (%)(k=2)
455 | 1098457 | ±1 | Power Linearity (%)
456 | 1098457 | ±1.5 | Spectral Compensation Accuracy (%)
457 | 1098457 | 514 | Calibration Wavelength (nanometers)
458 | 1098457 | Air | Cooling Method
459 | 1098457 | 2 | Cable Length (meters)
460 | 1098457 | None | Window
461 | 1098457 | None | Diffuser
462 | 1098457 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
463 | 1098457 | PM DB-25 | Cable Type
464 | 1098457 | 19 mm | Active Area Diameter
465 | 1098457 | 0.2 | Noise Equivalent Power (milliwatts)
466 | 1098457 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
467 | 1098461 | 600 | Maximum Energy Density (millijoules per square centimeter)
468 | 1098461 | 250 | Minimum Wavelength (nanometers)
469 | 1098461 | 11000 | Maximum Wavelength (nanometers)
470 | 1098461 | 100 | Minimum Power (watts)
471 | 1098461 | 5000 | Maximum Power (watts)
472 | 1098461 | Water | Cooling Type
473 | 1098461 | 0.25 to 11 | Wavelength Range (micrometers)
474 | 1098461 | 100W to 5000W | Power Range
475 | 1098461 | 7500 | Maximum Intermittent Power (watts)(<5 minutes)
476 | 1098461 | 1000 | Resolution (milliwatts)
477 | 1098461 | 1 to 2.5 | Maximum Power Density (kilowatts per square centimeter)
478 | 1098461 | Broadband | Detector Coating
479 | 1098461 | ±3 | Calibration Uncertainty (%)(k=2)
480 | 1098461 | ±1 | Power Linearity (%)
481 | 1098461 | ±1.5 | Spectral Compensation Accuracy (%)
482 | 1098461 | 1070 | Calibration Wavelength (nanometers)
483 | 1098461 | Water | Cooling Method
484 | 1098461 | 2 | Cable Length (meters)
485 | 1098461 | None | Window
486 | 1098461 | None | Diffuser
487 | 1098461 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
488 | 1098461 | PM DB-25 | Cable Type
489 | 1098461 | 100 mm | Active Area Diameter
490 | 1098461 | 20 | Noise Equivalent Power (milliwatts)
491 | 1098461 | 45 | Response Time, Speed-up On, 0% to 95% (seconds)
492 | 1098472 | 600 | Maximum Energy Density (millijoules per square centimeter)
493 | 1098472 | 250 | Minimum Wavelength (nanometers)
494 | 1098472 | 11000 | Maximum Wavelength (nanometers)
495 | 1098472 | 1 | Minimum Power (watts)
496 | 1098472 | 200 | Maximum Power (watts)
497 | 1098472 | 200 | Maximum Long-Pulse Energy (joules)
498 | 1098472 | Fan | Cooling Type
499 | 1098472 | 0.25 to 11 | Wavelength Range (micrometers)
500 | 1098472 | 1 W to 200W | Power Range
501 | 1098472 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
502 | 1098472 | 1 to 200 | Long-Pulse Joules Range (joules)
503 | 1098472 | 100 | Resolution (milliwatts)
504 | 1098472 | 6 | Maximum Power Density (kilowatts per square centimeter)
505 | 1098472 | Broadband | Detector Coating
506 | 1098472 | ±1 | Calibration Uncertainty (%)(k=2)
507 | 1098472 | ±1 | Power Linearity (%)
508 | 1098472 | ±1.5 | Spectral Compensation Accuracy (%)
509 | 1098472 | 514 | Calibration Wavelength (nanometers)
510 | 1098472 | Fan | Cooling Method
511 | 1098472 | 2 | Cable Length (meters)
512 | 1098472 | None | Window
513 | 1098472 | None | Diffuser
514 | 1098472 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
515 | 1098472 | PM DB-25 | Cable Type
516 | 1098472 | 50 mm | Active Area Diameter
517 | 1098472 | 5 mW (w/o fan), 100 mW (with fan) | Noise Equivalent Power (milliwatts)
518 | 1098472 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
519 | 1098480 | 600 | Maximum Energy Density (millijoules per square centimeter)
520 | 1098480 | 250 | Minimum Wavelength (nanometers)
521 | 1098480 | 11000 | Maximum Wavelength (nanometers)
522 | 1098480 | 1 | Minimum Power (watts)
523 | 1098480 | 200 | Maximum Power (watts)
524 | 1098480 | 200 | Maximum Long-Pulse Energy (joules)
525 | 1098480 | Fan | Cooling Type
526 | 1098480 | 0.25 to 11 | Wavelength Range (micrometers)
527 | 1098480 | 1 W to 200W | Power Range
528 | 1098480 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
529 | 1098480 | 1 to 200 | Long-Pulse Joules Range (joules)
530 | 1098480 | 100 | Resolution (milliwatts)
531 | 1098480 | 6 | Maximum Power Density (kilowatts per square centimeter)
532 | 1098480 | Broadband | Detector Coating
533 | 1098480 | ±1 | Calibration Uncertainty (%)(k=2)
534 | 1098480 | ±1 | Power Linearity (%)
535 | 1098480 | ±1.5 | Spectral Compensation Accuracy (%)
536 | 1098480 | 514 | Calibration Wavelength (nanometers)
537 | 1098480 | Fan | Cooling Method
538 | 1098480 | 2 | Cable Length (meters)
539 | 1098480 | None | Window
540 | 1098480 | None | Diffuser
541 | 1098480 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
542 | 1098480 | PM DB-25 | Cable Type
543 | 1098480 | 19 mm | Active Area Diameter
544 | 1098480 | 5 mW (w/o fan), 100 mW (with fan) | Noise Equivalent Power (milliwatts)
545 | 1098480 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
546 | 1098481 | 600 | Maximum Energy Density (millijoules per square centimeter)
547 | 1098481 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
548 | 1098481 | PM DB-25 | Cable Type
549 | 1098481 | 150 | Minimum Wavelength (nanometers)
550 | 1098481 | 1000 | Maximum Wavelength (nanometers)
551 | 1098481 | 1 | Minimum Power (watts)
552 | 1098481 | 300 | Maximum Power (watts)
553 | 1098481 | 300 | Maximum Long-Pulse Energy (joules)
554 | 1098481 | Fan | Cooling Type
555 | 1098481 | 0.15 to 1 | Wavelength Range (micrometers)
556 | 1098481 | 1W to 300W | Power Range
557 | 1098481 | 450 | Maximum Intermittent Power (watts)(<5 minutes)
558 | 1098481 | 1 to 300 | Long-Pulse Joules Range (joules)
559 | 1098481 | 100 | Resolution (milliwatts)
560 | 1098481 | 6 | Maximum Power Density (kilowatts per square centimeter)
561 | 1098481 | ±1 | Calibration Uncertainty (%)(k=2)
562 | 1098481 | ±1 | Power Linearity (%)
563 | 1098481 | ±1.5 | Spectral Compensation Accuracy (%)
564 | 1098481 | 514 | Calibration Wavelength (nanometers)
565 | 1098481 | Fan | Cooling Method
566 | 1098481 | 2 | Cable Length (meters)
567 | 1098481 | None | Window
568 | 1098481 | None | Diffuser
569 | 1098481 | 50 mm | Active Area Diameter
570 | 1098481 | 5 mW (w/o fan), 100 mW (with fan) | Noise Equivalent Power (milliwatts)
571 | 1098481 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
572 | 1098481 | UV | Detector Coating
573 | 1098490 | 600 | Maximum Energy Density (millijoules per square centimeter)
574 | 1098490 | 250 | Minimum Wavelength (nanometers)
575 | 1098490 | 11000 | Maximum Wavelength (nanometers)
576 | 1098490 | 100 | Minimum Power (watts)
577 | 1098490 | 1000 | Maximum Power (watts)
578 | 1098490 | Water | Cooling Type
579 | 1098490 | 0.25 to 11 | Wavelength Range (micrometers)
580 | 1098490 | 100W to 1000W | Power Range
581 | 1098490 | 1500 | Maximum Intermittent Power (watts)(<5 minutes)
582 | 1098490 | 1000 | Resolution (milliwatts)
583 | 1098490 | 1 to 2.5 | Maximum Power Density (kilowatts per square centimeter)
584 | 1098490 | Broadband | Detector Coating
585 | 1098490 | ±3 | Calibration Uncertainty (%)(k=2)
586 | 1098490 | ±1 | Power Linearity (%)
587 | 1098490 | ±1.5 | Spectral Compensation Accuracy (%)
588 | 1098490 | 1070 | Calibration Wavelength (nanometers)
589 | 1098490 | Water | Cooling Method
590 | 1098490 | 2 | Cable Length (meters)
591 | 1098490 | None | Window
592 | 1098490 | None | Diffuser
593 | 1098490 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
594 | 1098490 | PM DB-25 | Cable Type
595 | 1098490 | 100 mm | Active Area Diameter
596 | 1098490 | 20 | Noise Equivalent Power (milliwatts)
597 | 1098490 | 45 | Response Time, Speed-up On, 0% to 95% (seconds)
598 | 1098498 | 600 | Maximum Energy Density (millijoules per square centimeter)
599 | 1098498 | 150 | Minimum Wavelength (nanometers)
600 | 1098498 | 1000 | Maximum Wavelength (nanometers)
601 | 1098498 | 0.1 | Minimum Power (watts)
602 | 1098498 | 30 | Maximum Power (watts)
603 | 1098498 | 50 | Maximum Long-Pulse Energy (joules)
604 | 1098498 | Air | Cooling Type
605 | 1098498 | 0.15 - 1 | Wavelength Range (micrometers)
606 | 1098498 | 100 mW to 30W | Power Range
607 | 1098498 | 50 | Maximum Intermittent Power (watts)(<5 minutes)
608 | 1098498 | 0.5 to 50 | Long-Pulse Joules Range (joules)
609 | 1098498 | 10 | Resolution (milliwatts)
610 | 1098498 | 6 | Maximum Power Density (kilowatts per square centimeter)
611 | 1098498 | UV | Detector Coating
612 | 1098498 | ±1 | Calibration Uncertainty (%)(k=2)
613 | 1098498 | ±1 | Power Linearity (%)
614 | 1098498 | ±1.5 | Spectral Compensation Accuracy (%)
615 | 1098498 | 514 | Calibration Wavelength (nanometers)
616 | 1098498 | Air | Cooling Method
617 | 1098498 | 2 | Cable Length (meters)
618 | 1098498 | None | Window
619 | 1098498 | None | Diffuser
620 | 1098498 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
621 | 1098498 | PM DB-25 | Cable Type
622 | 1098498 | 19 mm | Active Area Diameter
623 | 1098498 | 0.5 | Noise Equivalent Power (milliwatts)
624 | 1098498 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
625 | 1098505 | 600 | Maximum Energy Density (millijoules per square centimeter)
626 | 1098505 | 250 | Minimum Wavelength (nanometers)
627 | 1098505 | 11000 | Maximum Wavelength (nanometers)
628 | 1098505 | 100 | Minimum Power (watts)
629 | 1098505 | 5000 | Maximum Power (watts)
630 | 1098505 | Water | Cooling Type
631 | 1098505 | 0.25 to 11 | Wavelength Range (micrometers)
632 | 1098505 | 100W to 5000W | Power Range
633 | 1098505 | 7500 | Maximum Intermittent Power (watts)(<5 minutes)
634 | 1098505 | 1000 | Resolution (milliwatts)
635 | 1098505 | 1 to 2.5 | Maximum Power Density (kilowatts per square centimeter)
636 | 1098505 | Broadband | Detector Coating
637 | 1098505 | ±3 | Calibration Uncertainty (%)(k=2)
638 | 1098505 | ±1 | Power Linearity (%)
639 | 1098505 | ±1.5 | Spectral Compensation Accuracy (%)
640 | 1098505 | 1070 | Calibration Wavelength (nanometers)
641 | 1098505 | Water | Cooling Method
642 | 1098505 | 2 | Cable Length (meters)
643 | 1098505 | None | Window
644 | 1098505 | None | Diffuser
645 | 1098505 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
646 | 1098505 | PM DB-25 | Cable Type
647 | 1098505 | 200 mm | Active Area Diameter
648 | 1098505 | 20 | Noise Equivalent Power (milliwatts)
649 | 1098505 | 45 | Response Time, Speed-up On, 0% to 95% (seconds)
650 | 1098506 | 600 | Maximum Energy Density (millijoules per square centimeter)
651 | 1098506 | 250 | Minimum Wavelength (nanometers)
652 | 1098506 | 11000 | Maximum Wavelength (nanometers)
653 | 1098506 | 100 | Minimum Power (watts)
654 | 1098506 | 3000 | Maximum Power (watts)
655 | 1098506 | Water | Cooling Type
656 | 1098506 | 0.25 to 11 | Wavelength Range (micrometers)
657 | 1098506 | 100W to 3000W | Power Range
658 | 1098506 | 4000 | Maximum Intermittent Power (watts)(<5 minutes)
659 | 1098506 | 1000 | Resolution (milliwatts)
660 | 1098506 | 1 to 2.5 | Maximum Power Density (kilowatts per square centimeter)
661 | 1098506 | Broadband | Detector Coating
662 | 1098506 | ±3 | Calibration Uncertainty (%)(k=2)
663 | 1098506 | ±1 | Power Linearity (%)
664 | 1098506 | ±1.5 | Spectral Compensation Accuracy (%)
665 | 1098506 | 1070 | Calibration Wavelength (nanometers)
666 | 1098506 | Water | Cooling Method
667 | 1098506 | 2 | Cable Length (meters)
668 | 1098506 | None | Window
669 | 1098506 | None | Diffuser
670 | 1098506 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
671 | 1098506 | PM DB-25 | Cable Type
672 | 1098506 | 100 mm | Active Area Diameter
673 | 1098506 | 20 | Noise Equivalent Power (milliwatts)
674 | 1098506 | 45 | Response Time, Speed-up On, 0% to 95% (seconds)
675 | 1098509 | 600 | Maximum Energy Density (millijoules per square centimeter)
676 | 1098509 | 250 | Minimum Wavelength (nanometers)
677 | 1098509 | 11000 | Maximum Wavelength (nanometers)
678 | 1098509 | 1 | Minimum Power (watts)
679 | 1098509 | 300 | Maximum Power (watts)
680 | 1098509 | 300 | Maximum Long-Pulse Energy (joules)
681 | 1098509 | Fan | Cooling Type
682 | 1098509 | 0.25 to 11 | Wavelength Range (micrometers)
683 | 1098509 | 1W to 300W | Power Range
684 | 1098509 | 450 | Maximum Intermittent Power (watts)(<5 minutes)
685 | 1098509 | 1 to 300 | Long-Pulse Joules Range (joules)
686 | 1098509 | 100 | Resolution (milliwatts)
687 | 1098509 | 6 | Maximum Power Density (kilowatts per square centimeter)
688 | 1098509 | Broadband | Detector Coating
689 | 1098509 | ±1 | Calibration Uncertainty (%)(k=2)
690 | 1098509 | ±1 | Power Linearity (%)
691 | 1098509 | ±1.5 | Spectral Compensation Accuracy (%)
692 | 1098509 | 514 | Calibration Wavelength (nanometers)
693 | 1098509 | Fan | Cooling Method
694 | 1098509 | 2 | Cable Length (meters)
695 | 1098509 | None | Window
696 | 1098509 | None | Diffuser
697 | 1098509 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
698 | 1098509 | PM DB-25 | Cable Type
699 | 1098509 | 19 mm | Active Area Diameter
700 | 1098509 | 5 mW (w/o fan), 100 mW (with fan) | Noise Equivalent Power (milliwatts)
701 | 1098509 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
702 | 1098579 | Average, Max, Min | Statistical Analysis
703 | 1098579 | 0 to 1, 2, or 5 VDC (selectable) | Analog Output Text
704 | 1098579 | 12VDC - center positive; (90 to 260 VAC, 50/60 Hz AC power adapter included) | Instrument Power
705 | 1098579 | CE, RoHS, WEEE, ISO 17025 | Compliance
707 | 1098579 | Optical | Sensor Compatibility
708 | 1098579 | Thermopile (PM-model) | Sensor Compatibility
709 | 1098579 | Power | Measurement Type
710 | 1098579 | Analog output | Outputs & Interfaces
711 | 1098579 | USB | Outputs & Interfaces
712 | 1098579 | Yes | Thermopile sensor compatibility
713 | 1098579 | Yes | Optical sensor compatibility
714 | 1098579 | No | Pyroelectric sensor compatibility
715 | 1098579 | No | PowerMax-Pro sensor compatibility
716 | 1098579 | Yes | USB Interface
717 | 1098579 | No | RS232 Interface
718 | 1098579 | Yes | Analog output
719 | 1098579 | No | External Trigger Input
720 | 1098579 | Fixed Segment LCD | Display Type
721 | 1098579 | Yes | Battery Operation
722 | 1098579 | Sensor dependent (reference sensor specifications) | Measurement Range
723 | 1098579 | 0.1% of full scale | Measurement Resolution
724 | 1098579 | N/A | Maximum Pulse Energy Repetition Rate (hertz)
725 | 1098579 | 10 Hz | Power Sampling Rate
726 | 1098579 | Meter accuracy + sensor accuracy | Accuracy - System
727 | 1098579 | ±1.0% | Accuracy - Analog Output
728 | 1098579 | ±1.0% | Calibration Uncertainty (%)(k=2)
729 | 1098579 | 58 x 73 mm fixed-segment LCD with backlight | Display
730 | 1098579 | USB, LabVIEW, PC application | Computer Interface
731 | 1098579 | N/A | Trigger
732 | 1098579 | 200 x 100 x 40 mm (7.87 x 3.94 x 1.57 in.) | Size (H x W x D)
733 | 1098579 | 1.0 kg (2.2. lbs.) | Weight
734 | 1098579 | Yes | C24
735 | 1098580 | Average, Max, Min, Std. Dev. | Statistical Analysis
736 | 1098580 | 0 to 1, 2, or 5 VDC (selectable) | Analog Output Text
737 | 1098580 | 12VDC - center positive; (90 to 260 VAC, 50/60 Hz AC power adapter included) | Instrument Power
738 | 1098580 | CE, RoHS, WEEE, ISO 17025 | Compliance
740 | 1098580 | 300 | Maximum Pyroelectric Rep. Rate (pulses per second)
741 | 1098580 | Optical | Sensor Compatibility
742 | 1098580 | Pyroelectric | Sensor Compatibility
743 | 1098580 | Thermopile (PM-model) | Sensor Compatibility
744 | 1098580 | Energy (pulse-to-pulse) | Measurement Type
745 | 1098580 | Energy (single-shot) | Measurement Type
746 | 1098580 | Power | Measurement Type
747 | 1098580 | Analog output | Outputs & Interfaces
748 | 1098580 | USB | Outputs & Interfaces
749 | 1098580 | Yes | Thermopile sensor compatibility
750 | 1098580 | Yes | Optical sensor compatibility
751 | 1098580 | Yes | Pyroelectric sensor compatibility
752 | 1098580 | No | PowerMax-Pro sensor compatibility
753 | 1098580 | Yes | USB Interface
754 | 1098580 | No | RS232 Interface
755 | 1098580 | Yes | Analog output
756 | 1098580 | No | External Trigger Input
757 | 1098580 | Fixed Segment LCD | Display Type
758 | 1098580 | Yes | Battery Operation
759 | 1098580 | Sensor dependent (reference sensor specifications) | Measurement Range
760 | 1098580 | 0.1% of full scale | Measurement Resolution
761 | 1098580 | 300 Hz (every pulse) | Maximum Pulse Energy Repetition Rate (hertz)
762 | 1098580 | 10 Hz | Power Sampling Rate
763 | 1098580 | Meter accuracy + sensor accuracy | Accuracy - System
764 | 1098580 | ±1.0% | Accuracy - Analog Output
765 | 1098580 | ±1.0% | Calibration Uncertainty (%)(k=2)
766 | 1098580 | 58 x 73 mm fixed-segment LCD with backlight | Display
767 | 1098580 | USB, LabVIEW, PC application | Computer Interface
768 | 1098580 | Internal trigger | Trigger
769 | 1098580 | 200 x 100 x 40 mm (7.87 x 3.94 x 1.57 in.) | Size (H x W x D)
770 | 1098580 | 1.0 kg (2.2. lbs.) | Weight
771 | 1098580 | Yes | C24
772 | 1110572 | 400 | Maximum Energy Density (millijoules per square centimeter)
773 | 1110572 | J DB-25 | Cable Type
774 | 1110572 | 100 to 999 pps | Maximum Repetition Rate
775 | 1110572 | 50 mm | Active Area Diameter
776 | 1110572 | 10 to 99 µsec | Maximum Laser Pulse Length
777 | 1110572 | 100 µsec to 1 msec | Maximum Laser Pulse Length
778 | 1110572 | 101 mJ to 1J | Maximum Pulse Energy
779 | 1110572 | 190 nm to 266 nm | Wavelength Range
780 | 1110572 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
781 | 1110572 | 86 | Maximum Pulse Width (microseconds)
782 | 1110572 | 2.5 | Cable Length (meters)
783 | 1110572 | 248 | Calibration Wavelength (nanometers)
784 | 1110572 | 500 µJ to 1J | Energy Range
785 | 1110572 | 520 | Maximum Energy Density (millijoules per square centimeter)
786 | 1110572 | 200 | Maximum Pyroelectric Rep. Rate (pulses per second)
787 | 1110572 | MaxUV | Detector Coating
788 | 1110572 | <16 µJ | Noise Equivalent Energy
789 | 1110572 | 15 | Maximum Average Power (watts)
790 | 1110572 | No | Diffuser
791 | 1110572 | ±3 | Calibration Uncertainty (%)(k=2)
792 | 1110572 | ±3 | Energy Linearity (%)
793 | 1110573 | 10 | Maximum Average Power (watts)
794 | 1110573 | 17 | Maximum Pulse Width (microseconds)
795 | 1110573 | 1 mJ - 2 J | Energy Range
796 | 1110573 | <33 µJ | Noise Equivalent Energy
797 | 1110573 | MaxBlack | Detector Coating
798 | 1110573 | No | Diffuser
799 | 1110573 | 1064 | Calibration Wavelength (nanometers)
800 | 1110573 | ±2 | Calibration Uncertainty (%)(k=2)
801 | 1110573 | ±3 | Energy Linearity (%)
802 | 1110573 | 2.5 | Cable Length (meters)
803 | 1110573 | 500 | Maximum Energy Density (millijoules per square centimeter)
804 | 1110573 | J DB-25 | Cable Type
805 | 1110573 | 100 to 999 pps | Maximum Repetition Rate
806 | 1110573 | 50 mm | Active Area Diameter
807 | 1110573 | 10 to 99 µsec | Maximum Laser Pulse Length
808 | 1110573 | Greater than 1J | Maximum Pulse Energy
809 | 1110573 | 190 nm to 12 µm | Wavelength Range
810 | 1110573 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
811 | 1110573 | 300 | Maximum Pyroelectric Rep. Rate (pulses per second)
812 | 1110574 | 20 | Maximum Average Power (watts)
813 | 1110574 | 1.7 | Maximum Pulse Width (microseconds)
814 | 1110574 | 500 µJ - 1 J | Energy Range
815 | 1110574 | <16 µJ | Noise Equivalent Energy
816 | 1110574 | Diffuse Metallic | Detector Coating
817 | 1110574 | No | Diffuser
818 | 1110574 | 1064 | Calibration Wavelength (nanometers)
819 | 1110574 | ±2 | Calibration Uncertainty (%)(k=2)
820 | 1110574 | ±3 | Energy Linearity (%)
821 | 1110574 | 2.5 | Cable Length (meters)
822 | 1110574 | 500 | Maximum Energy Density (millijoules per square centimeter)
823 | 1110574 | J DB-25 | Cable Type
824 | 1110574 | 10,000 pps | Maximum Repetition Rate
825 | 1110574 | 50 mm | Active Area Diameter
826 | 1110574 | Less than 10 µsec | Maximum Laser Pulse Length
827 | 1110574 | 101 mJ to 1J | Maximum Pulse Energy
828 | 1110574 | 190 nm to 2100 nm | Wavelength Range
829 | 1110574 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro | Meter Compatibility
830 | 1110575 | 200 | Maximum Energy Density (millijoules per square centimeter)
831 | 1110575 | J DB-25 | Cable Type
832 | 1110575 | 100 to 999 pps | Maximum Repetition Rate
833 | 1110575 | 50 mm | Active Area Diameter
834 | 1110575 | 10 to 99 µsec | Maximum Laser Pulse Length
835 | 1110575 | 100 µsec to 1 msec | Maximum Laser Pulse Length
836 | 1110575 | 101 mJ to 1J | Maximum Pulse Energy
837 | 1110575 | 190 nm to 266 nm | Wavelength Range
838 | 1110575 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
839 | 1110575 | 2.5 | Cable Length (meters)
840 | 1110575 | MaxUV | Detector Coating
841 | 1110575 | 86 | Maximum Pulse Width (microseconds)
842 | 1110575 | 125 µJ to 250 mJ | Energy Range
843 | 1110575 | <4 µJ | Noise Equivalent Energy
844 | 1110575 | 18 | Maximum Average Power
845 | 1110575 | 200 | Maximum Pyroelectric Rep. Rate (pulses per second)
846 | 1110575 | 400 | Maximum Energy Density (millijoules per square centimeter)
847 | 1110575 | DUV | Diffuser
848 | 1110575 | ±3 | Calibration Uncertainty (%)(k=2)
849 | 1110575 | ±3 | Energy Linearity (%)
850 | 1110575 | 193 | Calibration Wavelength (nanometers)
851 | 1110576 | 10 | Maximum Average Power (watts)
852 | 1110576 | 57 | Maximum Pulse Width (microseconds)
853 | 1110576 | 250 µJ - 500 mJ | Energy Range
854 | 1110576 | <8 µJ | Noise Equivalent Energy
855 | 1110576 | 0.19-12 | Wavelength Range (micrometers)
856 | 1110576 | MaxBlack | Detector Coating
857 | 1110576 | No | Diffuser
858 | 1110576 | 1064 | Calibration Wavelength (nanometers)
859 | 1110576 | ±2 | Calibration Uncertainty (%)(k=2)
860 | 1110576 | ±3 | Energy Linearity (%)
861 | 1110576 | 2.5 | Cable Length (meters)
862 | 1110576 | 500 | Maximum Energy Density (millijoules per square centimeter)
863 | 1110576 | J DB-25 | Cable Type
864 | 1110576 | 100 to 999 pps | Maximum Repetition Rate
865 | 1110576 | 50 mm | Active Area Diameter
866 | 1110576 | 10 to 99 µsec | Maximum Laser Pulse Length
867 | 1110576 | 101 mJ to 1J | Maximum Pulse Energy
868 | 1110576 | 190 nm to 12 µm | Wavelength Range
869 | 1110576 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
870 | 1110577 | 20 | Maximum Average Power (watts)
871 | 1110577 | 860 | Maximum Pulse Width (microseconds)
872 | 1110577 | 1.5 mJ - 3 J | Energy Range
873 | 1110577 | <50 µJ | Noise Equivalent Energy
874 | 1110577 | MaxBlack | Detector Coating
875 | 1110577 | IR | Diffuser
876 | 1110577 | 248 | Calibration Wavelength (nanometers)
877 | 1110577 | 5 | Maximum Energy Density (millijoules per square centimeter)
878 | 1110577 | J DB-25 | Cable Type
879 | 1110577 | Less than 100 pps | Maximum Repetition Rate
880 | 1110577 | 25 mm | Active Area Diameter
881 | 1110577 | 100 µsec to 1 msec | Maximum Laser Pulse Length
882 | 1110577 | Greater than 1J | Maximum Pulse Energy
883 | 1110577 | 532 nm to 2100 nm | Wavelength Range
884 | 1110577 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
885 | 1110577 | 20 | Maximum Pyroelectric Rep. Rate (pulses per second)
886 | 1110577 | ±2 | Calibration Uncertainty (%)(k=2)
887 | 1110577 | ±3 | Energy Linearity (%)
888 | 1110577 | 1064 | Calibration Wavelength (nanometers)
889 | 1110741 | J DB-25 | Cable Type
890 | 1110741 | 100 to 999 pps | Maximum Repetition Rate
891 | 1110741 | 25 mm | Active Area Diameter
892 | 1110741 | 10 to 99 µsec | Maximum Laser Pulse Length
893 | 1110741 | 1 mJ to 100 mJ | Maximum Pulse Energy
894 | 1110741 | 190 nm to 2100 nm | Wavelength Range
895 | 1110741 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
896 | 1110741 | MaxUV | Detector Coating
897 | 1110741 | 50 µJ to 100 mJ | Energy Range
898 | 1110741 | <2 µJ | Noise Equivalent Energy
899 | 1110741 | 5 | Maximum Average Power (watts)
900 | 1110741 | 43 | Maximum Pulse Width (microseconds)
901 | 1110741 | 500 | Maximum Pyroelectric Rep. Rate (pulses per second)
902 | 1110741 | ±3 | Energy Linearity (%)
903 | 1110741 | 200 | Maximum Energy Density (millijoules per square centimeter)
904 | 1110741 | No | Diffuser
905 | 1110741 | ±3 | Calibration Uncertainty (%)(k=2)
906 | 1110741 | 2.5 | Cable Length (meters)
907 | 1110741 | 193 | Calibration Wavelength (nanometers)
908 | 1110743 | 5 | Maximum Average Power (watts)
909 | 1110743 | 17 | Maximum Pulse Width (microseconds)
910 | 1110743 | 25 µJ - 50 mJ | Energy Range
911 | 1110743 | <1 µJ | Noise Equivalent Energy
912 | 1110743 | MaxBlack | Detector Coating
913 | 1110743 | No | Diffuser
914 | 1110743 | 1064 | Calibration Wavelength (nanometers)
915 | 1110743 | ±2 | Calibration Uncertainty (%)(k=2)
916 | 1110743 | ±3 | Energy Linearity (%)
917 | 1110743 | 2.5 | Cable Length (meters)
918 | 1110743 | 500 | Maximum Energy Density (millijoules per square centimeter)
919 | 1110743 | J DB-25 | Cable Type
920 | 1110743 | 1,000 pps | Maximum Repetition Rate
921 | 1110743 | 25 mm | Active Area Diameter
922 | 1110743 | 10 to 99 µsec | Maximum Laser Pulse Length
923 | 1110743 | 1 mJ to 100 mJ | Maximum Pulse Energy
924 | 1110743 | 190 nm to 12 µm | Wavelength Range
925 | 1110743 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
926 | 1110743 | 1000 | Maximum Pyroelectric Rep. Rate (pulses per second)
927 | 1110744 | 20 | Maximum Average Power (watts)
928 | 1110744 | 340 | Maximum Pulse Width (microseconds)
929 | 1110744 | 1.5 mJ - 3 J | Energy Range
930 | 1110744 | <50 µJ | Noise Equivalent Energy
931 | 1110744 | 0.266-2.1 | Wavelength Range (micrometers)
932 | 1110744 | 35 | Maximum Beam Size (millimeters)
933 | 1110744 | MaxBlack | Detector Coating
934 | 1110744 | YAG | Diffuser
935 | 1110744 | 1064 | Calibration Wavelength (nanometers)
936 | 1110744 | ±2 | Calibration Uncertainty (%)(k=2)
937 | 1110744 | ±3 | Energy Linearity (%)
938 | 1110744 | 2.5 | Cable Length (meters)
939 | 1110744 | 14000 (at 1064 nm, 10 ns) 2800 (at 532 nm, 10 ns) 750 (at 355 nm, 10 ns) 1000 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
940 | 1110744 | J DB-25 | Cable Type
941 | 1110744 | Less than 100 pps | Maximum Repetition Rate
942 | 1110744 | 50 mm | Active Area Diameter
943 | 1110744 | 100 µsec to 1 msec | Maximum Laser Pulse Length
944 | 1110744 | Greater than 1J | Maximum Pulse Energy
945 | 1110744 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
946 | 1110745 | J DB-25 | Cable Type
947 | 1110745 | 100 to 999 pps | Maximum Repetition Rate
948 | 1110745 | 25 mm | Active Area Diameter
949 | 1110745 | 10 to 99 µsec | Maximum Laser Pulse Length
950 | 1110745 | 101 mJ to 1J | Maximum Pulse Energy
951 | 1110745 | 190 nm to 2100 nm | Wavelength Range
952 | 1110745 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
953 | 1110745 | MaxUV | Detector Coating
954 | 1110745 | 125 µJ to 250 mJ | Energy Range
955 | 1110745 | <4 µJ | Noise Equivalent Energy
956 | 1110745 | 5 | Maximum Average Power (watts)
957 | 1110745 | 500 | Maximum Pyroelectric Rep. Rate (pulses per second)
958 | 1110745 | 248 | Calibration Wavelength (nanometers)
959 | 1110745 | 43 | Maximum Pulse Width (microseconds)
960 | 1110745 | 260 | Maximum Energy Density (millijoules per square centimeter)
961 | 1110745 | ±3 | Calibration Uncertainty (%)(k=2)
962 | 1110745 | No | Diffuser
963 | 1110745 | 2.5 | Cable Length (meters)
964 | 1110745 | ±3 | Energy Linearity (%)
965 | 1110746 | 5 | Maximum Average Power (watts)
966 | 1110746 | 17 | Maximum Pulse Width (microseconds)
967 | 1110746 | 500 µJ - 1 J | Energy Range
968 | 1110746 | <16 µJ | Noise Equivalent Energy
969 | 1110746 | MaxBlack | Detector Coating
970 | 1110746 | No | Diffuser
971 | 1110746 | 1064 | Calibration Wavelength (nanometers)
972 | 1110746 | ±2 | Calibration Uncertainty (%)(k=2)
973 | 1110746 | ±3 | Energy Linearity (%)
974 | 1110746 | 2.5 | Cable Length (meters)
975 | 1110746 | 500 | Maximum Energy Density (millijoules per square centimeter)
976 | 1110746 | J DB-25 | Cable Type
977 | 1110746 | 1,000 pps | Maximum Repetition Rate
978 | 1110746 | 25 mm | Active Area Diameter
979 | 1110746 | 10 to 99 µsec | Maximum Laser Pulse Length
980 | 1110746 | 101 mJ to 1J | Maximum Pulse Energy
981 | 1110746 | 190 nm to 12 µm | Wavelength Range
982 | 1110746 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
983 | 1110746 | 1000 | Maximum Pyroelectric Rep. Rate (pulses per second)
984 | 1110747 | 10 | Maximum Average Power (watts)
985 | 1110747 | 1.7 | Maximum Pulse Width (microseconds)
986 | 1110747 | 50 µJ - 100 mJ | Energy Range
987 | 1110747 | <2 µJ | Noise Equivalent Energy
988 | 1110747 | Diffuse Metallic | Detector Coating
989 | 1110747 | No | Diffuser
990 | 1110747 | 1064 | Calibration Wavelength (nanometers)
991 | 1110747 | ±2 | Calibration Uncertainty (%)(k=2)
992 | 1110747 | ±3 | Energy Linearity (%)
993 | 1110747 | 2.5 | Cable Length (meters)
994 | 1110747 | 500 | Maximum Energy Density (millijoules per square centimeter)
995 | 1110747 | J DB-25 | Cable Type
996 | 1110747 | 10,000 pps | Maximum Repetition Rate
997 | 1110747 | 25 mm | Active Area Diameter
998 | 1110747 | Less than 10 µsec | Maximum Laser Pulse Length
999 | 1110747 | 1 mJ to 100 mJ | Maximum Pulse Energy
1000 | 1110747 | 190 nm to 2100 nm | Wavelength Range
1001 | 1110747 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro | Meter Compatibility
1002 | 1110747 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1003 | 1110843 | 4 | Maximum Average Power (watts)
1004 | 1110843 | 17 | Maximum Pulse Width (microseconds)
1005 | 1110843 | 10 µJ - 20 mJ | Energy Range
1006 | 1110843 | <0.5 µJ | Noise Equivalent Energy
1007 | 1110843 | MaxBlack | Detector Coating
1008 | 1110843 | No | Diffuser
1009 | 1110843 | 1064 | Calibration Wavelength (nanometers)
1010 | 1110843 | ±2 | Calibration Uncertainty (%)(k=2)
1011 | 1110843 | ±3 | Energy Linearity (%)
1012 | 1110843 | 2.5 | Cable Length (meters)
1013 | 1110843 | 500 | Maximum Energy Density (millijoules per square centimeter)
1014 | 1110843 | J DB-25 | Cable Type
1015 | 1110843 | 1,000 pps | Maximum Repetition Rate
1016 | 1110843 | 10 mm | Active Area Diameter
1017 | 1110843 | 10 to 99 µsec | Maximum Laser Pulse Length
1018 | 1110843 | 1 mJ to 100 mJ | Maximum Pulse Energy
1019 | 1110843 | 190 nm to 12 µm | Wavelength Range
1020 | 1110843 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1021 | 1110843 | 1000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1022 | 1110855 | 4 | Maximum Average Power (watts)
1023 | 1110855 | 17 | Maximum Pulse Width (microseconds)
1024 | 1110855 | 300 nJ - 600 µJ | Energy Range
1025 | 1110855 | <20 nJ | Noise Equivalent Energy
1026 | 1110855 | MaxBlack | Detector Coating
1027 | 1110855 | No | Diffuser
1028 | 1110855 | 1064 | Calibration Wavelength (nanometers)
1029 | 1110855 | ±2 | Calibration Uncertainty (%)(k=2)
1030 | 1110855 | ±3 | Energy Linearity (%)
1031 | 1110855 | 2.5 | Cable Length (meters)
1032 | 1110855 | 500 | Maximum Energy Density (millijoules per square centimeter)
1033 | 1110855 | J DB-25 | Cable Type
1034 | 1110855 | 1,000 pps | Maximum Repetition Rate
1035 | 1110855 | 10 mm | Active Area Diameter
1036 | 1110855 | 10 to 99 µsec | Maximum Laser Pulse Length
1037 | 1110855 | 1 µJ to 1 mJ | Maximum Pulse Energy
1038 | 1110855 | 190 nm to 12 µm | Wavelength Range
1039 | 1110855 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1040 | 1110855 | 1000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1041 | 1110856 | 1 | Maximum Average Power (watts)
1042 | 1110856 | 1.7 | Maximum Pulse Width (microseconds)
1043 | 1110856 | 100 nJ - 200 µJ | Energy Range
1044 | 1110856 | <10 nJ | Noise Equivalent Energy
1045 | 1110856 | Diffuse Metallic | Detector Coating
1046 | 1110856 | 1064 | Calibration Wavelength (nanometers)
1047 | 1110856 | ±2 | Calibration Uncertainty (%)(k=2)
1048 | 1110856 | ±3 | Energy Linearity (%)
1049 | 1110856 | 2.5 | Cable Length (meters)
1050 | 1110856 | 50 | Maximum Energy Density (millijoules per square centimeter)
1051 | 1110856 | J DB-25 | Cable Type
1052 | 1110856 | 10,000 pps | Maximum Repetition Rate
1053 | 1110856 | 10 mm | Active Area Diameter
1054 | 1110856 | Less than 10 µsec | Maximum Laser Pulse Length
1055 | 1110856 | 1 µJ to 1 mJ | Maximum Pulse Energy
1056 | 1110856 | 190 nm to 2100 nm | Wavelength Range
1057 | 1110856 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro | Meter Compatibility
1058 | 1110856 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1059 | 1113493 | 600 | Maximum Energy Density (millijoules per square centimeter)
1060 | 1113493 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMate, FieldMaxII-TO, FieldMaxII-TOP | Meter Compatibility
1061 | 1113493 | 150 | Minimum Wavelength (nanometers)
1062 | 1113493 | 1000 | Maximum Wavelength (nanometers)
1063 | 1113493 | 1 | Minimum Power (watts)
1064 | 1113493 | 200 | Maximum Power (watts)
1065 | 1113493 | PM DB-25 | Cable Type
1066 | 1113493 | 200 | Maximum Long-Pulse Energy (joules)
1067 | 1113493 | Fan | Cooling Type
1068 | 1113493 | 0.15 to 1 | Wavelength Range (micrometers)
1069 | 1113493 | 1 W to 200W | Power Range
1070 | 1113493 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
1071 | 1113493 | 1 to 200 | Long-Pulse Joules Range (joules)
1072 | 1113493 | 100 | Resolution (milliwatts)
1073 | 1113493 | 6 | Maximum Power Density (kilowatts per square centimeter)
1074 | 1113493 | UV | Detector Coating
1075 | 1113493 | ±1 | Calibration Uncertainty (%)(k=2)
1076 | 1113493 | ±1 | Power Linearity (%)
1077 | 1113493 | ±1.5 | Spectral Compensation Accuracy (%)
1078 | 1113493 | 514 | Calibration Wavelength (nanometers)
1079 | 1113493 | Fan | Cooling Method
1080 | 1113493 | 2 | Cable Length (meters)
1081 | 1113493 | None | Window
1082 | 1113493 | None | Diffuser
1083 | 1113493 | 50 mm | Active Area Diameter
1084 | 1113493 | 5 mW (w/o fan), 100 mW (with fan) | Noise Equivalent Power (milliwatts)
1085 | 1113493 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
1086 | 1115484 | 600 | Maximum Energy Density (millijoules per square centimeter)
1087 | 1115484 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMate, FieldMaxII-TO, FieldMaxII-TOP | Meter Compatibility
1088 | 1115484 | PM DB-25 | Cable Type
1089 | 1115484 | 150 | Minimum Wavelength (nanometers)
1090 | 1115484 | 1000 | Maximum Wavelength (nanometers)
1091 | 1115484 | 1 | Minimum Power (watts)
1092 | 1115484 | 1000 | Maximum Power (watts)
1093 | 1115484 | Water | Cooling Type
1094 | 1115484 | 0.15 to 1 | Wavelength Range (micrometers)
1095 | 1115484 | 1W to 1000W | Power Range
1096 | 1115484 | 1500 | Maximum Intermittent Power (watts)(<5 minutes)
1097 | 1115484 | 1000 | Resolution (milliwatts)
1098 | 1115484 | 1 to 2.5 | Maximum Power Density (kilowatts per square centimeter)
1099 | 1115484 | 30 | Response Time, Speed-up On, 0% to 95% (seconds)
1100 | 1115484 | UV | Detector Coating
1101 | 1115484 | ±3 | Calibration Uncertainty (%)(k=2)
1102 | 1115484 | ±1 | Power Linearity (%)
1103 | 1115484 | ±1.5 | Spectral Compensation Accuracy (%)
1104 | 1115484 | 1070 | Calibration Wavelength (nanometers)
1105 | 1115484 | Water | Cooling Method
1106 | 1115484 | 2 | Cable Length (meters)
1107 | 1115484 | None | Window
1108 | 1115484 | None | Diffuser
1109 | 1115484 | 50 mm | Active Area Diameter
1110 | 1115484 | 20 | Noise Equivalent Power (milliwatts)
1111 | 1140408 | J DB-25 | Cable Type
1112 | 1140408 | 10,000 pps | Maximum Repetition Rate
1113 | 1140408 | 10 mm | Active Area Diameter
1114 | 1140408 | Less than 10 µsec | Maximum Laser Pulse Length
1115 | 1140408 | Less than 1 µJ | Maximum Pulse Energy
1116 | 1140408 | 800 nm to 1700 nm | Wavelength Range
1117 | 1140408 | Germanium | Detector
1118 | 1140408 | 200 pJ - 600 nJ | Energy Range
1119 | 1140408 | <8 pJ (at 1064 nm) | Noise Equivalent Energy
1120 | 1140408 | 15 | Maximum Average Power (watts)
1121 | 1140408 | 1064 | Calibration Wavelength (nanometers)
1122 | 1140408 | 1 | Maximum Pulse Width (microseconds)
1123 | 1140408 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1124 | 1140408 | No | Diffuser
1125 | 1140408 | ±3 | Calibration Uncertainty (%)(k=2)
1126 | 1140408 | 3 | Cable Length (meters)
1127 | 1140727 | J DB-25 | Cable Type
1128 | 1140727 | 10,000 pps | Maximum Repetition Rate
1129 | 1140727 | 10 mm | Active Area Diameter
1130 | 1140727 | Less than 10 µsec | Maximum Laser Pulse Length
1131 | 1140727 | Less than 1 µJ | Maximum Pulse Energy
1132 | 1140727 | 325 nm to 900 nm | Wavelength Range
1133 | 1140727 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1134 | 1140727 | 0.8 pJ (at 532 nm) | Noise Equivalent Energy
1135 | 1140727 | ±3 | Calibration Uncertainty (%)(k=2)
1136 | 1140727 | Silicon | Detector
1137 | 1140727 | 6 mW | Maximum Average Power
1138 | 1140727 | 1 | Maximum Pulse Width (microseconds)
1139 | 1140727 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1140 | 1140727 | 3 | Cable Length (meters)
1141 | 1140727 | ND2 | Diffuser
1142 | 1140727 | ±3 | Energy Linearity (%)
1143 | 1140727 | 532 | Calibration Wavelength (nanometers)
1144 | 1141474 | 600 | Maximum Energy Density (millijoules per square centimeter)
1145 | 1141474 | 250 | Minimum Wavelength (nanometers)
1146 | 1141474 | 11000 | Maximum Wavelength (nanometers)
1147 | 1141474 | 1 | Minimum Power (watts)
1148 | 1141474 | 300 | Maximum Power (watts)
1149 | 1141474 | Water | Cooling Type
1150 | 1141474 | 0.25 - 11 | Wavelength Range (micrometers)
1151 | 1141474 | 1 W to 300W | Power Range
1152 | 1141474 | 450 | Maximum Intermittent Power (watts)(<5 minutes)
1153 | 1141474 | 6 | Maximum Power Density (kilowatts per square centimeter)
1154 | 1141474 | Broadband | Detector Coating
1155 | 1141474 | ±1 | Calibration Uncertainty (%)(k=2)
1156 | 1141474 | ±1 | Power Linearity (%)
1157 | 1141474 | ±1.5 | Spectral Compensation Accuracy (%)
1158 | 1141474 | 514 | Calibration Wavelength (nanometers)
1159 | 1141474 | Water | Cooling Method
1160 | 1141474 | 2 | Cable Length (meters)
1161 | 1141474 | None | Window
1162 | 1141474 | None | Diffuser
1163 | 1141474 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMate, FieldMaxII-TO, FieldMaxII-TOP | Meter Compatibility
1164 | 1141474 | PM DB-25 | Cable Type
1165 | 1141474 | 19 mm | Active Area Diameter
1166 | 1141474 | 20 | Noise Equivalent Power (milliwatts)
1167 | 1141474 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
1168 | 1144701 | 20 | Maximum Average Power (watts)
1169 | 1144701 | 57 | Maximum Pulse Width (microseconds)
1170 | 1144701 | 1.5 mJ - 3 J | Energy Range
1171 | 1144701 | <50 µJ | Noise Equivalent Energy
1172 | 1144701 | 35 | Maximum Beam Size (millimeters)
1173 | 1144701 | MaxBlack | Detector Coating
1174 | 1144701 | YAG | Diffuser
1175 | 1144701 | 1064 | Calibration Wavelength (nanometers)
1176 | 1144701 | ±2 | Calibration Uncertainty (%)(k=2)
1177 | 1144701 | ±3 | Energy Linearity (%)
1178 | 1144701 | 2.5 | Cable Length (meters)
1179 | 1144701 | 14000 (at 1064 nm, 10 ns) 2800 (at 532 nm, 10 ns) 750 (at 355 nm, 10 ns) 1000 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
1180 | 1144701 | J DB-25 | Cable Type
1181 | 1144701 | 100 to 999 pps | Maximum Repetition Rate
1182 | 1144701 | 50 mm | Active Area Diameter
1183 | 1144701 | 10 to 99 µsec | Maximum Laser Pulse Length
1184 | 1144701 | Greater than 1J | Maximum Pulse Energy
1185 | 1144701 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1186 | 1146237 | 260 | Maximum Energy Density (millijoules per square centimeter)
1187 | 1146237 | J DB-25 | Cable Type
1188 | 1146237 | 100 to 999 pps | Maximum Repetition Rate
1189 | 1146237 | 50 mm | Active Area Diameter
1190 | 1146237 | 10 to 99 µsec | Maximum Laser Pulse Length
1191 | 1146237 | 100 µsec to 1 msec | Maximum Laser Pulse Length
1192 | 1146237 | 101 mJ to 1J | Maximum Pulse Energy
1193 | 1146237 | 190 nm to 2100 nm | Wavelength Range
1194 | 1146237 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1195 | 1146237 | MaxUV | Detector Coating
1196 | 1146237 | 125 µJ to 250 mJ | Energy Range
1197 | 1146237 | <4 µJ | Noise Equivalent Energy
1198 | 1146237 | 10 | Maximum Average Power (watts)
1199 | 1146237 | 86 | Maximum Pulse Width (microseconds)
1200 | 1146237 | 200 | Maximum Pyroelectric Rep. Rate (pulses per second)
1201 | 1146237 | 200 | Maximum Energy Density (millijoules per square centimeter)
1202 | 1146237 | DUV | Diffuser
1203 | 1146237 | ±3 | Calibration Uncertainty (%)(k=2)
1204 | 1146237 | ±3 | Energy Linearity (%)
1205 | 1146237 | 193 | Calibration Wavelength (nanometers)
1206 | 1146237 | 2.5 | Cable Length (meters)
1207 | 1146243 | 200 | Maximum Energy Density (millijoules per square centimeter)
1208 | 1146243 | J DB-25 | Cable Type
1209 | 1146243 | 100 to 999 pps | Maximum Repetition Rate
1210 | 1146243 | 50 mm | Active Area Diameter
1211 | 1146243 | 10 to 99 µsec | Maximum Laser Pulse Length
1212 | 1146243 | 100 µsec to 1 msec | Maximum Laser Pulse Length
1213 | 1146243 | 101 mJ to 1J | Maximum Pulse Energy
1214 | 1146243 | 190 nm to 2100 nm | Wavelength Range
1215 | 1146243 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1216 | 1146243 | ±3 | Energy Linearity (%)
1217 | 1146243 | 2.5 | Cable Length (meters)
1218 | 1146243 | 500 µJ to 1J | Energy Range
1219 | 1146243 | 86 | Maximum Pulse Width (microseconds)
1220 | 1146243 | No | Diffuser
1221 | 1146243 | MaxUV | Detector Coating
1222 | 1146243 | <16 µJ | Noise Equivalent Energy
1223 | 1146243 | 10 | Maximum Average Power (watts)
1224 | 1146243 | 200 | Maximum Pyroelectric Rep. Rate (pulses per second)
1225 | 1146243 | 260 | Maximum Energy Density (millijoules per square centimeter)
1226 | 1146243 | ±3 | Calibration Uncertainty (%)(k=2)
1227 | 1146243 | 248 | Calibration Wavelength (nanometers)
1228 | 1150146 | J DB-25 | Cable Type
1229 | 1150146 | 10,000 pps | Maximum Repetition Rate
1230 | 1150146 | 10 mm | Active Area Diameter
1231 | 1150146 | Less than 10 µsec | Maximum Laser Pulse Length
1232 | 1150146 | Less than 1 µJ | Maximum Pulse Energy
1233 | 1150146 | 325 nm to 900 nm | Wavelength Range
1234 | 1150146 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1235 | 1150146 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1236 | 1151431 | 20 | Maximum Average Power (watts)
1237 | 1151431 | 2000 | Maximum Pulse Width (microseconds)
1238 | 1151431 | 12 mJ to 15J | Energy Range
1239 | 1151431 | <50 µJ | Noise Equivalent Energy
1240 | 1151431 | 30 | Maximum Beam Size (millimeters)
1241 | 1151431 | MaxBlack | Detector Coating
1242 | 1151431 | IR | Diffuser
1243 | 1151431 | 1064 & 2940 | Calibration Wavelength (nanometers)
1244 | 1151431 | ±2 (1064nm); ±3 (2940nm) | Calibration Uncertainty (%)(k=2)
1245 | 1151431 | ±3.5 | Energy Linearity (%)
1246 | 1151431 | 2.5 | Cable Length (meters)
1247 | 1151431 | 14 (at 1064 nm, 10 ns) 2.8 (at 532 nm, 10 ns) 1 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
1248 | 1151431 | J DB-25 | Cable Type
1249 | 1151431 | Less than 100 pps | Maximum Repetition Rate
1250 | 1151431 | 50 mm | Active Area Diameter
1251 | 1151431 | Greater than 1 msec | Maximum Laser Pulse Length
1252 | 1151431 | Greater than 1 msec | Maximum Laser Pulse Length
1253 | 1151431 | Greater than 1J | Maximum Pulse Energy
1254 | 1151431 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1255 | 1152086 | 600 | Maximum Energy Density (millijoules per square centimeter)
1256 | 1152086 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMate, FieldMaxII-TO, FieldMaxII-TOP | Meter Compatibility
1257 | 1152086 | PM DB-25 | Cable Type
1258 | 1152086 | 150 | Minimum Wavelength (nanometers)
1259 | 1152086 | 1000 | Maximum Wavelength (nanometers)
1260 | 1152086 | 1 | Minimum Power (watts)
1261 | 1152086 | 1000 | Maximum Power (watts)
1262 | 1152086 | Water | Cooling Type
1263 | 1152086 | 0.15 to 1 | Wavelength Range (micrometers)
1264 | 1152086 | 1W to 1000W | Power Range
1265 | 1152086 | 1500 | Maximum Intermittent Power (watts)(<5 minutes)
1266 | 1152086 | 1000 | Resolution (milliwatts)
1267 | 1152086 | 1 to 2.5 | Maximum Power Density (kilowatts per square centimeter)
1268 | 1152086 | UV | Detector Coating
1269 | 1152086 | ±3 | Calibration Uncertainty (%)(k=2)
1270 | 1152086 | ±1 | Power Linearity (%)
1271 | 1152086 | ±1.5 | Spectral Compensation Accuracy (%)
1272 | 1152086 | 1070 | Calibration Wavelength (nanometers)
1273 | 1152086 | Water | Cooling Method
1274 | 1152086 | 2 | Cable Length (meters)
1275 | 1152086 | None | Window
1276 | 1152086 | None | Diffuser
1277 | 1152086 | 100 mm | Active Area Diameter
1278 | 1152086 | 20 | Noise Equivalent Power (milliwatts)
1279 | 1152086 | 45 | Response Time, Speed-up On, 0% to 95% (seconds)
1280 | 1155722 | 15 | Maximum Average Power (watts)
1281 | 1155722 | 1000 | Maximum Pulse Width (microseconds)
1282 | 1155722 | 1 mJ-3J | Energy Range
1283 | 1155722 | <100 µJ | Noise Equivalent Energy
1284 | 1155722 | MaxBlack | Detector Coating
1285 | 1155722 | No | Diffuser
1286 | 1155722 | 248 | Calibration Wavelength (nanometers)
1287 | 1155722 | ±3 | Calibration Uncertainty (%)(k=2)
1288 | 1155722 | ±3 | Energy Linearity (%)
1289 | 1155722 | 2.5 | Cable Length (meters)
1290 | 1155722 | 260 | Maximum Energy Density (millijoules per square centimeter)
1291 | 1155722 | J DB-25 | Cable Type
1292 | 1155722 | Less than 100 pps | Maximum Repetition Rate
1293 | 1155722 | 50 mm | Active Area Diameter
1294 | 1155722 | 100 µsec to 1 msec | Maximum Laser Pulse Length
1295 | 1155722 | Greater than 1J | Maximum Pulse Energy
1296 | 1155722 | 500 nm to 3 µm | Wavelength Range
1297 | 1155722 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMaxII-TOP | Meter Compatibility
1298 | 1155722 | 35 | Maximum Beam Size (millimeters)
1299 | 1168337 | 325 | Minimum Wavelength (nanometers)
1300 | 1168337 | 1065 | Maximum Wavelength (nanometers)
1301 | 1168337 | 10000 | Minimum Power (nanowatts)
1302 | 1168337 | 100 | Max Power (milliwatts)
1303 | 1168337 | 325 nm to 1065 nm | Wavelength Range (micrometers)
1304 | 1168337 | 10 µW to >100 mW (wavelength dependent) | Power Range
1305 | 1168337 | 100 | Noise Equivalent Power (nanowatts)
1306 | 1168337 | 20 | Maximum Power Density (watts per square centimeter)
1307 | 1168337 | Silicon | Detector
1308 | 1168337 | ±1 | Calibration Uncertainty (%)(k=2)
1309 | 1168337 | ±1 | Power Linearity (%)
1310 | 1168337 | ±4 (325 to 900 nm)/±5 (900 to 1065 nm) | Spectral Compensation Accuracy (%)
1311 | 1168337 | 514 | Calibration Wavelength (nanometers)
1312 | 1168337 | Air | Cooling Method
1313 | 1168337 | 2.5 | Cable Length (meters)
1314 | 1168337 | None | Window
1315 | 1168337 | Yes | Diffuser
1316 | 1168337 | USB | Cable Type
1317 | 1168337 | 100 mW | Maximum Power
1318 | 1168337 | 5 µW | Minimum Power
1319 | 1168337 | 325 nm to 1065 nm | Wavelength Range
1320 | 1168337 | 10 mm | Active Area Diameter
1321 | 1168337 | Yes | C24
1322 | 1168343 | 50 | Maximum Energy Density (millijoules per square centimeter)
1323 | 1168343 | USB | Cable Type
1324 | 1168343 | 300 | Minimum Wavelength (nanometers)
1325 | 1168343 | 2100 | Maximum Wavelength (nanometers)
1326 | 1168343 | 0.0001 | Minimum Power (watts)
1327 | 1168343 | 1 | Maximum Power (watts)
1328 | 1168343 | 1 | Maximum Long-Pulse Energy (joules)
1329 | 1168343 | Air | Cooling Type
1330 | 1168343 | 0.3 to 2.1 µm | Wavelength Range (micrometers)
1331 | 1168343 | 100 µW to 1W | Power Range
1332 | 1168343 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
1333 | 1168343 | 0.001 to 1 | Long-Pulse Joules Range (joules)
1334 | 1168343 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
1335 | 1168343 | ±1 | Calibration Uncertainty (%)(k=2)
1336 | 1168343 | ±1 | Power Linearity (%)
1337 | 1168343 | ±1.5 | Spectral Compensation Accuracy (%)
1338 | 1168343 | 514 | Calibration Wavelength (nanometers)
1339 | 1168343 | Air | Cooling Method
1340 | 1168343 | 2.5 | Cable Length (meters)
1341 | 1168343 | Yes | Window
1342 | 1168343 | None | Diffuser
1343 | 1168343 | 19 mm | Active Area Diameter
1344 | 1168343 | 0.003 | Noise Equivalent Power (milliwatts)
1345 | 1168343 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1346 | 1168343 | Black | Detector Coating
1347 | 1168344 | Water, Air (intermittent) | Cooling Type
1348 | 1168344 | 5, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
1349 | 1168344 | 600 | Maximum Energy Density (millijoules per square centimeter)
1350 | 1168344 | Water, Air (intermittent) | Cooling Method
1351 | 1168344 | USB | Cable Type
1352 | 1168344 | 190 | Minimum Wavelength (nanometers)
1353 | 1168344 | 11000 | Maximum Wavelength (nanometers)
1354 | 1168344 | 0.01 | Minimum Power (watts)
1355 | 1168344 | 10 | Maximum Power (watts)
1356 | 1168344 | 10 | Maximum Long-Pulse Energy (joules)
1357 | 1168344 | 0.19 to 11 µm | Wavelength Range (micrometers)
1358 | 1168344 | 10mW to 10W | Power Range
1359 | 1168344 | 0.5 to 10 | Long-Pulse Joules Range (joules)
1360 | 1168344 | 0.2 | Noise Equivalent Power (milliwatts)
1361 | 1168344 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1362 | 1168344 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1363 | 1168344 | Broadband | Detector Coating
1364 | 1168344 | ±2 | Calibration Uncertainty (%)(k=2)
1365 | 1168344 | ±1 | Power Linearity (%)
1366 | 1168344 | ±1.5 | Spectral Compensation Accuracy (%)
1367 | 1168344 | 10600 | Calibration Wavelength (nanometers)
1368 | 1168344 | 2.5 | Cable Length (meters)
1369 | 1168344 | None | Window
1370 | 1168344 | None | Diffuser
1371 | 1168344 | 19 mm | Active Area Diameter
1372 | 1168345 | Water, Air (intermittent) | Cooling Type
1373 | 1168345 | 5, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
1374 | 1168345 | 600 | Maximum Energy Density (millijoules per square centimeter)
1375 | 1168345 | Water, Air (intermittent) | Cooling Method
1376 | 1168345 | RS232 | Cable Type
1377 | 1168345 | 190 | Minimum Wavelength (nanometers)
1378 | 1168345 | 11000 | Maximum Wavelength (nanometers)
1379 | 1168345 | 0.01 | Minimum Power (watts)
1380 | 1168345 | 10 | Maximum Power (watts)
1381 | 1168345 | 10 | Maximum Long-Pulse Energy (joules)
1382 | 1168345 | 0.19 to 11 µm | Wavelength Range (micrometers)
1383 | 1168345 | 10 mW to 10W (water-cooled) | Power Range
1384 | 1168345 | 0.5 to 10 | Long-Pulse Joules Range (joules)
1385 | 1168345 | 0.2 | Noise Equivalent Power (milliwatts)
1386 | 1168345 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1387 | 1168345 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1388 | 1168345 | Broadband | Detector Coating
1389 | 1168345 | ±2 | Calibration Uncertainty (%)(k=2)
1390 | 1168345 | ±1 | Power Linearity (%)
1391 | 1168345 | ±1.5 | Spectral Compensation Accuracy (%)
1392 | 1168345 | 10600 | Calibration Wavelength (nanometers)
1393 | 1168345 | 0.3 | Cable Length (meters)
1394 | 1168345 | None | Window
1395 | 1168345 | None | Diffuser
1396 | 1168345 | 19 mm | Active Area Diameter
1397 | 1168346 | Water, Air (intermittent) | Cooling Type
1398 | 1168346 | 20, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
1399 | 1168346 | 600 | Maximum Energy Density (millijoules per square centimeter)
1400 | 1168346 | Water, Air (intermittent) | Cooling Method
1401 | 1168346 | USB | Cable Type
1402 | 1168346 | 190 | Minimum Wavelength (nanometers)
1403 | 1168346 | 11000 | Maximum Wavelength (nanometers)
1404 | 1168346 | 0.3 | Minimum Power (watts)
1405 | 1168346 | 150 | Maximum Power (watts)
1406 | 1168346 | 150 | Maximum Long-Pulse Energy (joules)
1407 | 1168346 | 0.19 to 11 µm | Wavelength Range (micrometers)
1408 | 1168346 | 300 mW to 150W (water-cooled) | Power Range
1409 | 1168346 | 1 to 150 | Long-Pulse Joules Range (joules)
1410 | 1168346 | 1 | Noise Equivalent Power (milliwatts)
1411 | 1168346 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1412 | 1168346 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1413 | 1168346 | Broadband | Detector Coating
1414 | 1168346 | ±2 | Calibration Uncertainty (%)(k=2)
1415 | 1168346 | ±1 | Power Linearity (%)
1416 | 1168346 | ±1.5 | Spectral Compensation Accuracy (%)
1417 | 1168346 | 10600 | Calibration Wavelength (nanometers)
1418 | 1168346 | 2.5 | Cable Length (meters)
1419 | 1168346 | None | Window
1420 | 1168346 | None | Diffuser
1421 | 1168346 | 19 mm | Active Area Diameter
1422 | 1168347 | Water, Air (intermittent) | Cooling Type
1423 | 1168347 | 20, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
1424 | 1168347 | 600 | Maximum Energy Density (millijoules per square centimeter)
1425 | 1168347 | Water, Air (intermittent) | Cooling Method
1426 | 1168347 | RS232 | Cable Type
1427 | 1168347 | 190 | Minimum Wavelength (nanometers)
1428 | 1168347 | 11000 | Maximum Wavelength (nanometers)
1429 | 1168347 | 0.3 | Minimum Power (watts)
1430 | 1168347 | 150 | Maximum Power (watts)
1431 | 1168347 | 150 | Maximum Long-Pulse Energy (joules)
1432 | 1168347 | 0.19 to 11 µm | Wavelength Range (micrometers)
1433 | 1168347 | 300 mW to 150W (water-cooled) | Power Range
1434 | 1168347 | 1 to 150 | Long-Pulse Joules Range (joules)
1435 | 1168347 | 1 | Noise Equivalent Power (milliwatts)
1436 | 1168347 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1437 | 1168347 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1438 | 1168347 | Broadband | Detector Coating
1439 | 1168347 | ±2 | Calibration Uncertainty (%)(k=2)
1440 | 1168347 | ±1 | Power Linearity (%)
1441 | 1168347 | ±1.5 | Spectral Compensation Accuracy (%)
1442 | 1168347 | 10600 | Calibration Wavelength (nanometers)
1443 | 1168347 | 0.3 | Cable Length (meters)
1444 | 1168347 | None | Window
1445 | 1168347 | None | Diffuser
1446 | 1168347 | 19 mm | Active Area Diameter
1447 | 1168348 | Water, Air (intermittent) | Cooling Type
1448 | 1168348 | 80, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
1449 | 1168348 | 600 | Maximum Energy Density (millijoules per square centimeter)
1450 | 1168348 | Water, Air (intermittent) | Cooling Method
1451 | 1168348 | USB | Cable Type
1452 | 1168348 | 190 | Minimum Wavelength (nanometers)
1453 | 1168348 | 11000 | Maximum Wavelength (nanometers)
1454 | 1168348 | 0.3 | Minimum Power (watts)
1455 | 1168348 | 150 | Maximum Power (watts)
1456 | 1168348 | 150 | Maximum Long-Pulse Energy (joules)
1457 | 1168348 | 0.19 to 11 µm | Wavelength Range (micrometers)
1458 | 1168348 | 300 mW to 150W (water-cooled) | Power Range
1459 | 1168348 | 1 to 150 | Long-Pulse Joules Range (joules)
1460 | 1168348 | 1 | Noise Equivalent Power (milliwatts)
1461 | 1168348 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1462 | 1168348 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
1463 | 1168348 | Broadband | Detector Coating
1464 | 1168348 | ±2 | Calibration Uncertainty (%)(k=2)
1465 | 1168348 | ±1 | Power Linearity (%)
1466 | 1168348 | ±1.5 | Spectral Compensation Accuracy (%)
1467 | 1168348 | 10600 | Calibration Wavelength (nanometers)
1468 | 1168348 | 2.5 | Cable Length (meters)
1469 | 1168348 | None | Window
1470 | 1168348 | None | Diffuser
1471 | 1168348 | 50 mm | Active Area Diameter
1472 | 1168349 | 600 | Maximum Energy Density (millijoules per square centimeter)
1473 | 1168349 | Water, Air (intermittent) | Cooling Method
1474 | 1168349 | Water, Air (intermittent) | Cooling Type
1475 | 1168349 | 80, air-cooled | Maximum Intermittent Power (watts)(<5 minutes)
1476 | 1168349 | RS232 | Cable Type
1477 | 1168349 | 190 | Minimum Wavelength (nanometers)
1478 | 1168349 | 11000 | Maximum Wavelength (nanometers)
1479 | 1168349 | 0.3 | Minimum Power (watts)
1480 | 1168349 | 150 | Maximum Power (watts)
1481 | 1168349 | 150 | Maximum Long-Pulse Energy (joules)
1482 | 1168349 | 0.19 to 11 µm | Wavelength Range (micrometers)
1483 | 1168349 | 300 mW to 150W (water-cooled | Power Range
1484 | 1168349 | 1 to 150 | Long-Pulse Joules Range (joules)
1485 | 1168349 | 1 | Noise Equivalent Power (milliwatts)
1486 | 1168349 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1487 | 1168349 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
1488 | 1168349 | Broadband | Detector Coating
1489 | 1168349 | ±2 | Calibration Uncertainty (%)(k=2)
1490 | 1168349 | ±1 | Power Linearity (%)
1491 | 1168349 | ±1.5 | Spectral Compensation Accuracy (%)
1492 | 1168349 | 10600 | Calibration Wavelength (nanometers)
1493 | 1168349 | 0.3 | Cable Length (meters)
1494 | 1168349 | None | Window
1495 | 1168349 | None | Diffuser
1496 | 1168349 | 50 mm | Active Area Diameter
1497 | 1174257 | 600 | Maximum Energy Density (millijoules per square centimeter)
1498 | 1174257 | USB | Cable Type
1499 | 1174257 | 190 | Minimum Wavelength (nanometers)
1500 | 1174257 | 11000 | Maximum Wavelength (nanometers)
1501 | 1174257 | 0.01 | Minimum Power (watts)
1502 | 1174257 | 30 | Maximum Power (watts)
1503 | 1174257 | 50 | Maximum Long-Pulse Energy (joules)
1504 | 1174257 | Air | Cooling Type
1505 | 1174257 | 0.19 to 11 µm | Wavelength Range (micrometers)
1506 | 1174257 | 10 mW to 30W | Power Range
1507 | 1174257 | 50 | Maximum Intermittent Power (watts)(<5 minutes)
1508 | 1174257 | 0.5 to 50 | Long-Pulse Joules Range (joules)
1509 | 1174257 | 0.5 | Noise Equivalent Power (milliwatts)
1510 | 1174257 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1511 | 1174257 | Broadband | Detector Coating
1512 | 1174257 | ±2 | Calibration Uncertainty (%)(k=2)
1513 | 1174257 | ±1 | Power Linearity (%)
1514 | 1174257 | ±1.5 | Spectral Compensation Accuracy (%)
1515 | 1174257 | 10600 | Calibration Wavelength (nanometers)
1516 | 1174257 | Air | Cooling Method
1517 | 1174257 | 2.5 | Cable Length (meters)
1518 | 1174257 | None | Window
1519 | 1174257 | None | Diffuser
1520 | 1174257 | 19 mm | Active Area Diameter
1521 | 1174257 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
1522 | 1174258 | 600 | Maximum Energy Density (millijoules per square centimeter)
1523 | 1174258 | RS232 | Cable Type
1524 | 1174258 | 190 | Minimum Wavelength (nanometers)
1525 | 1174258 | 11000 | Maximum Wavelength (nanometers)
1526 | 1174258 | 0.01 | Minimum Power (watts)
1527 | 1174258 | 30 | Maximum Power (watts)
1528 | 1174258 | 50 | Maximum Long-Pulse Energy (joules)
1529 | 1174258 | Air | Cooling Type
1530 | 1174258 | 0.19 to 11 µm | Wavelength Range (micrometers)
1531 | 1174258 | 10 mW to 30W | Power Range
1532 | 1174258 | 50 | Maximum Intermittent Power (watts)(<5 minutes)
1533 | 1174258 | 0.5 to 50 | Long-Pulse Joules Range (joules)
1534 | 1174258 | 0.5 | Noise Equivalent Power (milliwatts)
1535 | 1174258 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1536 | 1174258 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
1537 | 1174258 | ±2 | Calibration Uncertainty (%)(k=2)
1538 | 1174258 | ±1 | Power Linearity (%)
1539 | 1174258 | ±1.5 | Spectral Compensation Accuracy (%)
1540 | 1174258 | 10600 | Calibration Wavelength (nanometers)
1541 | 1174258 | Air | Cooling Method
1542 | 1174258 | 0.3 | Cable Length (meters)
1543 | 1174258 | None | Window
1544 | 1174258 | None | Diffuser
1545 | 1174258 | 19 mm | Active Area Diameter
1546 | 1174258 | Broadband | Detector Coating
1547 | 1174259 | 600 | Maximum Energy Density (millijoules per square centimeter)
1548 | 1174259 | RS232 | Cable Type
1549 | 1174259 | 150 | Minimum Wavelength (nanometers)
1550 | 1174259 | 1000 | Maximum Wavelength (nanometers)
1551 | 1174259 | 0.01 | Minimum Power (watts)
1552 | 1174259 | 30 | Maximum Power (watts)
1553 | 1174259 | 50 | Maximum Long-Pulse Energy (joules)
1554 | 1174259 | Air | Cooling Type
1555 | 1174259 | 0.15 to 1 µm | Wavelength Range (micrometers)
1556 | 1174259 | 10 mW to 30W | Power Range
1557 | 1174259 | 50 | Maximum Intermittent Power (watts)(<5 minutes)
1558 | 1174259 | 0.5 to 50 | Long-Pulse Joules Range (joules)
1559 | 1174259 | 0.5 | Noise Equivalent Power (milliwatts)
1560 | 1174259 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1561 | 1174259 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
1562 | 1174259 | UV | Detector Coating
1563 | 1174259 | ±1 | Calibration Uncertainty (%)(k=2)
1564 | 1174259 | ±1 | Power Linearity (%)
1565 | 1174259 | ±1.5 | Spectral Compensation Accuracy (%)
1566 | 1174259 | 514 | Calibration Wavelength (nanometers)
1567 | 1174259 | Air | Cooling Method
1568 | 1174259 | 0.3 | Cable Length (meters)
1569 | 1174259 | None | Window
1570 | 1174259 | None | Diffuser
1571 | 1174259 | 19 mm | Active Area Diameter
1572 | 1174260 | 50 | Maximum Energy Density (millijoules per square centimeter)
1573 | 1174260 | USB | Cable Type
1574 | 1174260 | 300 | Minimum Wavelength (nanometers)
1575 | 1174260 | 11000 | Maximum Wavelength (nanometers)
1576 | 1174260 | 0.0001 | Minimum Power (watts)
1577 | 1174260 | 1 | Maximum Power (watts)
1578 | 1174260 | 1 | Maximum Long-Pulse Energy (joules)
1579 | 1174260 | Air | Cooling Type
1580 | 1174260 | 0.19 to 11 µm | Wavelength Range (micrometers)
1581 | 1174260 | 100 uW to 1W | Power Range
1582 | 1174260 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
1583 | 1174260 | 0.001 to 1 | Long-Pulse Joules Range (joules)
1584 | 1174260 | 0.003 | Noise Equivalent Power (milliwatts)
1585 | 1174260 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
1586 | 1174260 | ±2 | Calibration Uncertainty (%)(k=2)
1587 | 1174260 | ±1 | Power Linearity (%)
1588 | 1174260 | ±1.5 | Spectral Compensation Accuracy (%)
1589 | 1174260 | 514 | Calibration Wavelength (nanometers)
1590 | 1174260 | Air | Cooling Method
1591 | 1174260 | 2.5 | Cable Length (meters)
1592 | 1174260 | None | Window
1593 | 1174260 | None | Diffuser
1594 | 1174260 | 10 mm | Active Area Diameter
1595 | 1174260 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1596 | 1174260 | Black | Detector Coating
1597 | 1174261 | 50 | Maximum Energy Density (millijoules per square centimeter)
1598 | 1174261 | USB | Cable Type
1599 | 1174261 | 300 | Minimum Wavelength (nanometers)
1600 | 1174261 | 11000 | Maximum Wavelength (nanometers)
1601 | 1174261 | 0.0001 | Minimum Power (watts)
1602 | 1174261 | 1 | Maximum Power (watts)
1603 | 1174261 | 1 | Maximum Long-Pulse Energy (joules)
1604 | 1174261 | Air | Cooling Type
1605 | 1174261 | 0.19 to 11 µm | Wavelength Range (micrometers)
1606 | 1174261 | 100 uW to 1W | Power Range
1607 | 1174261 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
1608 | 1174261 | 0.001 to 1 | Long-Pulse Joules Range (joules)
1609 | 1174261 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
1610 | 1174261 | ±2 | Calibration Uncertainty (%)(k=2)
1611 | 1174261 | ±1 | Power Linearity (%)
1612 | 1174261 | ±1.5 | Spectral Compensation Accuracy (%)
1613 | 1174261 | 514 | Calibration Wavelength (nanometers)
1614 | 1174261 | Air | Cooling Method
1615 | 1174261 | 2.5 | Cable Length (meters)
1616 | 1174261 | None | Window
1617 | 1174261 | None | Diffuser
1618 | 1174261 | 19 mm | Active Area Diameter
1619 | 1174261 | 0.005 | Noise Equivalent Power (milliwatts)
1620 | 1174261 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1621 | 1174261 | Black | Detector Coating
1622 | 1174262 | 600 | Maximum Energy Density (millijoules per square centimeter)
1623 | 1174262 | USB | Cable Type
1624 | 1174262 | 190 | Minimum Wavelength (nanometers)
1625 | 1174262 | 11000 | Maximum Wavelength (nanometers)
1626 | 1174262 | 0.005 | Minimum Power (watts)
1627 | 1174262 | 10 | Maximum Power (watts)
1628 | 1174262 | 10 | Maximum Long-Pulse Energy (joules)
1629 | 1174262 | Air | Cooling Type
1630 | 1174262 | 0.19 to 11 µm | Wavelength Range (micrometers)
1631 | 1174262 | 5 mW to 10W | Power Range
1632 | 1174262 | 30 | Maximum Intermittent Power (watts)(<5 minutes)
1633 | 1174262 | 0.5 to 10 | Long-Pulse Joules Range (joules)
1634 | 1174262 | 0.2 | Noise Equivalent Power (milliwatts)
1635 | 1174262 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1636 | 1174262 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1637 | 1174262 | Broadband | Detector Coating
1638 | 1174262 | ±2 | Calibration Uncertainty (%)(k=2)
1639 | 1174262 | ±1 | Power Linearity (%)
1640 | 1174262 | ±1.5 | Spectral Compensation Accuracy (%)
1641 | 1174262 | 10600 | Calibration Wavelength (nanometers)
1642 | 1174262 | Air | Cooling Method
1643 | 1174262 | 2.5 | Cable Length (meters)
1644 | 1174262 | None | Window
1645 | 1174262 | None | Diffuser
1646 | 1174262 | 19 mm | Active Area Diameter
1647 | 1174263 | 50 | Maximum Energy Density (millijoules per square centimeter)
1648 | 1174263 | USB | Cable Type
1649 | 1174263 | 300 | Minimum Wavelength (nanometers)
1650 | 1174263 | 11000 | Maximum Wavelength (nanometers)
1651 | 1174263 | 0.0005 | Minimum Power (watts)
1652 | 1174263 | 2 | Maximum Power (watts)
1653 | 1174263 | Air | Cooling Type
1654 | 1174263 | 0.3 to 11 µm | Wavelength Range (micrometers)
1655 | 1174263 | 500 uW to 2W | Power Range
1656 | 1174263 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
1657 | 1174263 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
1658 | 1174263 | ±2 | Calibration Uncertainty (%)(k=2)
1659 | 1174263 | ±1 | Power Linearity (%)
1660 | 1174263 | ±1.5 | Spectral Compensation Accuracy (%)
1661 | 1174263 | 10600 | Calibration Wavelength (nanometers)
1662 | 1174263 | Air | Cooling Method
1663 | 1174263 | 2.5 | Cable Length (meters)
1664 | 1174263 | None | Window
1665 | 1174263 | None | Diffuser
1666 | 1174263 | 19 mm | Active Area Diameter
1667 | 1174263 | 0.02 | Noise Equivalent Power (milliwatts)
1668 | 1174263 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1669 | 1174263 | Black | Detector Coating
1670 | 1174264 | 600 | Maximum Energy Density (millijoules per square centimeter)
1671 | 1174264 | USB | Cable Type
1672 | 1174264 | 190 | Minimum Wavelength (nanometers)
1673 | 1174264 | 11000 | Maximum Wavelength (nanometers)
1674 | 1174264 | 0.005 | Minimum Power (watts)
1675 | 1174264 | 2 | Maximum Power (watts)
1676 | 1174264 | 2 | Maximum Long-Pulse Energy (joules)
1677 | 1174264 | Air | Cooling Type
1678 | 1174264 | 0.19 to 11 µm | Wavelength Range (micrometers)
1679 | 1174264 | 5 mW to 2W | Power Range
1680 | 1174264 | 5 | Maximum Intermittent Power (watts)(<5 minutes)
1681 | 1174264 | 0.5 to 2 | Long-Pulse Joules Range (joules)
1682 | 1174264 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1683 | 1174264 | ±2 | Calibration Uncertainty (%)(k=2)
1684 | 1174264 | ±1 | Power Linearity (%)
1685 | 1174264 | ±1.5 | Spectral Compensation Accuracy (%)
1686 | 1174264 | 10600 | Calibration Wavelength (nanometers)
1687 | 1174264 | Air | Cooling Method
1688 | 1174264 | 2.5 | Cable Length (meters)
1689 | 1174264 | None | Window
1690 | 1174264 | None | Diffuser
1691 | 1174264 | 19 mm | Active Area Diameter
1692 | 1174264 | 0.2 | Noise Equivalent Power (milliwatts)
1693 | 1174264 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1694 | 1174264 | Broadband | Detector Coating
1695 | 1174267 | 500 | Maximum Energy Density (millijoules per square centimeter)
1696 | 1174267 | RS232 | Cable Type
1697 | 1174267 | 250 | Minimum Wavelength (nanometers)
1698 | 1174267 | 10600 | Maximum Wavelength (nanometers)
1699 | 1174267 | 50 | Minimum Power (watts)
1700 | 1174267 | 1000 | Maximum Power (watts)
1701 | 1174267 | Water | Cooling Type
1702 | 1174267 | 0.25 to 10.6 µm | Wavelength Range (micrometers)
1703 | 1174267 | 50 to 1000W | Power Range
1704 | 1174267 | 20 | Noise Equivalent Power (milliwatts)
1705 | 1174267 | 1 to 2.5 kW/cm² - power dependent | Maximum Power Density (kilowatts per square centimeter)
1706 | 1174267 | Broadband | Detector Coating
1707 | 1174267 | ±5 | Calibration Uncertainty (%)(k=2)
1708 | 1174267 | ±1 | Power Linearity (%)
1709 | 1174267 | ±1.5 | Spectral Compensation Accuracy (%)
1710 | 1174267 | 10600 | Calibration Wavelength (nanometers)
1711 | 1174267 | Water | Cooling Method
1712 | 1174267 | 0.3 | Cable Length (meters)
1713 | 1174267 | None | Window
1714 | 1174267 | None | Diffuser
1715 | 1174267 | 36 mm | Active Area Diameter
1716 | 1174267 | 4 | Response Time, Speed-up On, 0% to 95% (seconds)
1717 | 1174756 | 20 | Maximum Average Power (watts)
1718 | 1174756 | 340 | Maximum Pulse Width (microseconds)
1719 | 1174756 | 50 µJ to 100 mJ | Energy Range
1720 | 1174756 | <50 µJ | Noise Equivalent Energy
1721 | 1174756 | 35 | Maximum Beam Size (millimeters)
1722 | 1174756 | MaxBlack | Detector Coating
1723 | 1174756 | IR | Diffuser
1724 | 1174756 | 1064 | Calibration Wavelength (nanometers)
1725 | 1174756 | ±2 | Calibration Uncertainty (%)(k=2)
1726 | 1174756 | ±3 | Energy Linearity (%)
1727 | 1174756 | 2.5 | Cable Length (meters)
1728 | 1174756 | 14 (at 1064 nm, 10 ns) 2.8 (at 532 nm, 10 ns) 1 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
1729 | 1174756 | J DB-25 | Cable Type
1730 | 1174756 | Less than 100 pps | Maximum Repetition Rate
1731 | 1174756 | 50 mm | Active Area Diameter
1732 | 1174756 | 100 µsec to 1 msec | Maximum Laser Pulse Length
1733 | 1174756 | 1 mJ to 100 mJ | Maximum Pulse Energy
1734 | 1179504 | RS232 | Cable Type
1735 | 1179504 | 300 | Minimum Wavelength (nanometers)
1736 | 1179504 | 2100 | Maximum Wavelength (nanometers)
1737 | 1179504 | 0.0001 | Minimum Power (watts)
1738 | 1179504 | 1 | Maximum Power (watts)
1739 | 1179504 | 1 | Maximum Long-Pulse Energy (joules)
1740 | 1179504 | Air | Cooling Type
1741 | 1179504 | 0.3 to 2.1 µm | Wavelength Range (micrometers)
1742 | 1179504 | 100 uW to 1W | Power Range
1743 | 1179504 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
1744 | 1179504 | 0.001 to 1 | Long-Pulse Joules Range (joules)
1745 | 1179504 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
1746 | 1179504 | -- | Maximum Energy Density (millijoules per square centimeter)
1747 | 1179504 | ±1 | Calibration Uncertainty (%)(k=2)
1748 | 1179504 | ±1 | Power Linearity (%)
1749 | 1179504 | ±1.5 | Spectral Compensation Accuracy (%)
1750 | 1179504 | 514 | Calibration Wavelength (nanometers)
1751 | 1179504 | Air | Cooling Method
1752 | 1179504 | 0.3 | Cable Length (meters)
1753 | 1179504 | Yes | Window
1754 | 1179504 | None | Diffuser
1755 | 1179504 | 19 mm | Active Area Diameter
1756 | 1179504 | 0.003 | Noise Equivalent Power (milliwatts)
1757 | 1179504 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1758 | 1179504 | Black | Detector Coating
1759 | 1184874 | 600 | Maximum Energy Density (millijoules per square centimeter)
1760 | 1184874 | USB | Cable Type
1761 | 1184874 | 150 | Minimum Wavelength (nanometers)
1762 | 1184874 | 1000 | Maximum Wavelength (nanometers)
1763 | 1184874 | 0.005 | Minimum Power (watts)
1764 | 1184874 | 10 | Maximum Power (watts)
1765 | 1184874 | 10 | Maximum Long-Pulse Energy (joules)
1766 | 1184874 | Air | Cooling Type
1767 | 1184874 | 0.15 to 1 µm | Wavelength Range (micrometers)
1768 | 1184874 | 5 mW to 10W | Power Range
1769 | 1184874 | 30 | Maximum Intermittent Power (watts)(<5 minutes)
1770 | 1184874 | 0.5 to 10 | Long-Pulse Joules Range (joules)
1771 | 1184874 | 0.2 | Noise Equivalent Power (milliwatts)
1772 | 1184874 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
1773 | 1184874 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
1774 | 1184874 | UV | Detector Coating
1775 | 1184874 | ±1 | Calibration Uncertainty (%)(k=2)
1776 | 1184874 | ±1 | Power Linearity (%)
1777 | 1184874 | ±1.5 | Spectral Compensation Accuracy (%)
1778 | 1184874 | 514 | Calibration Wavelength (nanometers)
1779 | 1184874 | Air | Cooling Method
1780 | 1184874 | 2.5 | Cable Length (meters)
1781 | 1184874 | None | Window
1782 | 1184874 | None | Diffuser
1783 | 1184874 | 19 mm | Active Area Diameter
1784 | 1191427 | 0.06 | Maximum Average Power (watts)
1785 | 1191427 | 1 | Maximum Pulse Width (microseconds)
1786 | 1191427 | 750 pJ to 775 nJ (at 532 nm) | Energy Range
1787 | 1191427 | <75 pJ (at 532 nm) | Noise Equivalent Energy
1788 | 1191427 | .325 to .900 | Wavelength Range (micrometers)
1789 | 1191427 | ND2 | Diffuser
1790 | 1191427 | 1064 | Calibration Wavelength (nanometers)
1791 | 1191427 | ±3 | Calibration Uncertainty (%)(k=2)
1792 | 1191427 | ±3 | Energy Linearity (%)
1793 | 1191427 | 3 | Cable Length (meters)
1794 | 1191427 | RS232 | Cable Type
1795 | 1191427 | 10,000 pps | Maximum Repetition Rate
1796 | 1191427 | 10 mm | Active Area Diameter
1797 | 1191427 | Less than 10 µsec | Maximum Laser Pulse Length
1798 | 1191427 | Less than 1 µJ | Maximum Pulse Energy
1799 | 1191427 | 325 nm to 900 nm | Wavelength Range
1800 | 1191427 | Silicon | Detector
1801 | 1191427 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1802 | 1191428 | 4 | Maximum Average Power (watts)
1803 | 1191428 | 17 | Maximum Pulse Width (microseconds)
1804 | 1191428 | 500 nJ to 600 µJ | Energy Range
1805 | 1191428 | <50 nJ | Noise Equivalent Energy
1806 | 1191428 | 0.19 to 12 | Wavelength Range (micrometers)
1807 | 1191428 | MaxBlack | Detector Coating
1808 | 1191428 | No | Diffuser
1809 | 1191428 | 1064 | Calibration Wavelength (nanometers)
1810 | 1191428 | ±2 | Calibration Uncertainty (%)(k=2)
1811 | 1191428 | ±3 | Energy Linearity (%)
1812 | 1191428 | -- | Cable Length (meters)
1813 | 1191428 | 500 | Maximum Energy Density (millijoules per square centimeter)
1814 | 1191428 | RS232 | Cable Type
1815 | 1191428 | 1,000 pps | Maximum Repetition Rate
1816 | 1191428 | 10 mm | Active Area Diameter
1817 | 1191428 | 10 to 99 µsec | Maximum Laser Pulse Length
1818 | 1191428 | 1 µJ to 1 mJ | Maximum Pulse Energy
1819 | 1191428 | 190 nm to 12 µm | Wavelength Range
1820 | 1191428 | 1000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1821 | 1191429 | 4 | Maximum Average Power (watts)
1822 | 1191429 | 17 | Maximum Pulse Width (microseconds)
1823 | 1191429 | 12 µJ to 20 mJ | Energy Range
1824 | 1191429 | <1.2 µJ | Noise Equivalent Energy
1825 | 1191429 | 0.19 to 12 | Wavelength Range (micrometers)
1826 | 1191429 | MaxBlack | Detector Coating
1827 | 1191429 | No | Diffuser
1828 | 1191429 | 1064 | Calibration Wavelength (nanometers)
1829 | 1191429 | ±2 | Calibration Uncertainty (%)(k=2)
1830 | 1191429 | ±3 | Energy Linearity (%)
1831 | 1191429 | 3 | Cable Length (meters)
1832 | 1191429 | 500 | Maximum Energy Density (millijoules per square centimeter)
1833 | 1191429 | RS232 | Cable Type
1834 | 1191429 | 1,000 pps | Maximum Repetition Rate
1835 | 1191429 | 10 mm | Active Area Diameter
1836 | 1191429 | 10 to 99 µsec | Maximum Laser Pulse Length
1837 | 1191429 | 1 mJ to 100 mJ | Maximum Pulse Energy
1838 | 1191429 | 190 nm to 12 µm | Wavelength Range
1839 | 1191430 | 20 | Maximum Average Power (watts)
1840 | 1191430 | 340 | Maximum Pulse Width (microseconds)
1841 | 1191430 | 2.4 mJ to 3J | Energy Range
1842 | 1191430 | <240 µJ | Noise Equivalent Energy
1843 | 1191430 | 0.266 to 2.1 | Wavelength Range (micrometers)
1844 | 1191430 | 35 | Maximum Beam Size (millimeters)
1845 | 1191430 | MaxBlack | Detector Coating
1846 | 1191430 | YAG | Diffuser
1847 | 1191430 | ±3 | Energy Linearity (%)
1848 | 1191430 | 3 | Cable Length (meters)
1849 | 1191430 | 14 | Maximum Energy Density (millijoules per square centimeter)
1850 | 1191430 | 1064 | Calibration Wavelength (nanometers)
1851 | 1191430 | ±2 | Calibration Uncertainty (%)(k=2)
1852 | 1191430 | RS232 | Cable Type
1853 | 1191430 | Less than 100 pps | Maximum Repetition Rate
1854 | 1191430 | 100 µsec to 1 msec | Maximum Laser Pulse Length
1855 | 1191430 | Greater than 1J | Maximum Pulse Energy
1856 | 1191430 | 50 | Maximum Pyroelectric Rep. Rate (pulses per second)
1857 | 1191431 | 5 | Maximum Average Power (watts)
1858 | 1191431 | 17 | Maximum Pulse Width (microseconds)
1859 | 1191431 | 50 µJ to 50 mJ | Energy Range
1860 | 1191431 | <5 µJ | Noise Equivalent Energy
1861 | 1191431 | 0.19 to 12 | Wavelength Range (micrometers)
1862 | 1191431 | MaxBlack | Detector Coating
1863 | 1191431 | No | Diffuser
1864 | 1191431 | 1064 | Calibration Wavelength (nanometers)
1865 | 1191431 | ±2 | Calibration Uncertainty (%)(k=2)
1866 | 1191431 | ±3 | Energy Linearity (%)
1867 | 1191431 | 3 | Cable Length (meters)
1868 | 1191431 | 500 | Maximum Energy Density (millijoules per square centimeter)
1869 | 1191431 | RS232 | Cable Type
1870 | 1191431 | 1,000 pps | Maximum Repetition Rate
1871 | 1191431 | 25 mm | Active Area Diameter
1872 | 1191431 | 10 to 99 µsec | Maximum Laser Pulse Length
1873 | 1191431 | 1 mJ to 100 mJ | Maximum Pulse Energy
1874 | 1191431 | 190 nm to 12 µm | Wavelength Range
1875 | 1191432 | 10 | Maximum Average Power (watts)
1876 | 1191432 | 57 | Maximum Pulse Width (microseconds)
1877 | 1191432 | 1.6 mJ to 2J | Energy Range
1878 | 1191432 | <160 µJ | Noise Equivalent Energy
1879 | 1191432 | 0.19 to 12 | Wavelength Range (micrometers)
1880 | 1191432 | MaxBlack | Detector Coating
1881 | 1191432 | No | Diffuser
1882 | 1191432 | 1064 | Calibration Wavelength (nanometers)
1883 | 1191432 | ±2 | Calibration Uncertainty (%)(k=2)
1884 | 1191432 | ±3 | Energy Linearity (%)
1885 | 1191432 | 3 | Cable Length (meters)
1886 | 1191432 | 500 | Maximum Energy Density (millijoules per square centimeter)
1887 | 1191432 | RS232 | Cable Type
1888 | 1191432 | 100 to 999 pps | Maximum Repetition Rate
1889 | 1191432 | 50 mm | Active Area Diameter
1890 | 1191432 | 10 to 99 µsec | Maximum Laser Pulse Length
1891 | 1191432 | Greater than 1J | Maximum Pulse Energy
1892 | 1191432 | 190 nm to 12 µm | Wavelength Range
1893 | 1191432 | 300 | Maximum Pyroelectric Rep. Rate (pulses per second)
1894 | 1191433 | 20 | Maximum Average Power (watts)
1895 | 1191433 | 1.7 | Maximum Pulse Width (microseconds)
1896 | 1191433 | 400 µJ to 1J | Energy Range
1897 | 1191433 | <40 µJ | Noise Equivalent Energy
1898 | 1191433 | 0.19 to 2.1 | Wavelength Range (micrometers)
1899 | 1191433 | Diffuse Metallic | Detector Coating
1900 | 1191433 | No | Diffuser
1901 | 1191433 | 1064 | Calibration Wavelength (nanometers)
1902 | 1191433 | ±2 | Calibration Uncertainty (%)(k=2)
1903 | 1191433 | ±3 | Energy Linearity (%)
1904 | 1191433 | 3 | Cable Length (meters)
1905 | 1191433 | 500 | Maximum Energy Density (millijoules per square centimeter)
1906 | 1191433 | RS232 | Cable Type
1907 | 1191433 | 10,000 pps | Maximum Repetition Rate
1908 | 1191433 | 50 mm | Active Area Diameter
1909 | 1191433 | Less than 10 µsec | Maximum Laser Pulse Length
1910 | 1191433 | 101 mJ to 1J | Maximum Pulse Energy
1911 | 1191433 | 190 nm to 2100 nm | Wavelength Range
1912 | 1191433 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
1913 | 1191434 | 0.06 | Maximum Average Power (watts)
1914 | 1191434 | 1 | Maximum Pulse Width (microseconds)
1915 | 1191434 | 750 pJ to 775 nJ (at 532 nm) | Energy Range
1916 | 1191434 | <75 pJ (at 532 nm) | Noise Equivalent Energy
1917 | 1191434 | .325 to .900 | Wavelength Range (micrometers)
1918 | 1191434 | Silicon | Detector Coating
1919 | 1191434 | ND2 | Diffuser
1920 | 1191434 | 532 | Calibration Wavelength (nanometers)
1921 | 1191434 | ±3 | Calibration Uncertainty (%)(k=2)
1922 | 1191434 | ±3 | Energy Linearity (%)
1923 | 1191434 | 3 | Cable Length (meters)
1924 | 1191434 | USB | Cable Type
1925 | 1191434 | 10,000 pps | Maximum Repetition Rate
1926 | 1191434 | 10 mm | Active Area Diameter
1927 | 1191434 | Less than 10 µsec | Maximum Laser Pulse Length
1928 | 1191434 | Less than 1 µJ | Maximum Pulse Energy
1929 | 1191434 | 325 nm to 900 nm | Wavelength Range
1930 | 1191435 | 4 | Maximum Average Power (watts)
1931 | 1191435 | 17 | Maximum Pulse Width (microseconds)
1932 | 1191435 | 500 nJ to 600 µJ | Energy Range
1933 | 1191435 | <50 nJ | Noise Equivalent Energy
1934 | 1191435 | 0.19 to 12 | Wavelength Range (micrometers)
1935 | 1191435 | MaxBlack | Detector Coating
1936 | 1191435 | No | Diffuser
1937 | 1191435 | 1064 | Calibration Wavelength (nanometers)
1938 | 1191435 | ±2 | Calibration Uncertainty (%)(k=2)
1939 | 1191435 | ±3 | Energy Linearity (%)
1940 | 1191435 | 3 | Cable Length (meters)
1941 | 1191435 | 500 | Maximum Energy Density (millijoules per square centimeter)
1942 | 1191435 | USB | Cable Type
1943 | 1191435 | 1,000 pps | Maximum Repetition Rate
1944 | 1191435 | 10 mm | Active Area Diameter
1945 | 1191435 | 10 to 99 µsec | Maximum Laser Pulse Length
1946 | 1191435 | 1 µJ to 1 mJ | Maximum Pulse Energy
1947 | 1191435 | 190 nm to 12 µm | Wavelength Range
1948 | 1191436 | 4 | Maximum Average Power (watts)
1949 | 1191436 | 17 | Maximum Pulse Width (microseconds)
1950 | 1191436 | 12 µJ to 20 mJ | Energy Range
1951 | 1191436 | <1.2 µJ | Noise Equivalent Energy
1952 | 1191436 | 0.19 to 12 | Wavelength Range (micrometers)
1953 | 1191436 | MaxBlack | Detector Coating
1954 | 1191436 | No | Diffuser
1955 | 1191436 | 1064 | Calibration Wavelength (nanometers)
1956 | 1191436 | ±2 | Calibration Uncertainty (%)(k=2)
1957 | 1191436 | ±3 | Energy Linearity (%)
1958 | 1191436 | 3 | Cable Length (meters)
1959 | 1191436 | 500 | Maximum Energy Density (millijoules per square centimeter)
1960 | 1191436 | USB | Cable Type
1961 | 1191436 | 1,000 pps | Maximum Repetition Rate
1962 | 1191436 | 10 mm | Active Area Diameter
1963 | 1191436 | 10 to 99 µsec | Maximum Laser Pulse Length
1964 | 1191436 | 1 mJ to 100 mJ | Maximum Pulse Energy
1965 | 1191436 | 190 nm to 12 µm | Wavelength Range
1966 | 1191437 | 20 | Maximum Average Power (watts)
1967 | 1191437 | 340 | Maximum Pulse Width (microseconds)
1968 | 1191437 | 2.4 mJ to 3J | Energy Range
1969 | 1191437 | <240 µJ | Noise Equivalent Energy
1970 | 1191437 | 0.266 to 2.1 | Wavelength Range (micrometers)
1971 | 1191437 | 35 | Maximum Beam Size (millimeters)
1972 | 1191437 | MaxBlack | Detector Coating
1973 | 1191437 | YAG | Diffuser
1974 | 1191437 | 1064 | Calibration Wavelength (nanometers)
1975 | 1191437 | ±2 | Calibration Uncertainty (%)(k=2)
1976 | 1191437 | ±3 | Energy Linearity (%)
1977 | 1191437 | 3 | Cable Length (meters)
1978 | 1191437 | 14000 (at 1064 nm, 10 ns) 2800 (at 532 nm, 10 ns) 750 (at 355 nm, 10 ns) 1000 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
1979 | 1191437 | USB | Cable Type
1980 | 1191437 | Less than 100 pps | Maximum Repetition Rate
1981 | 1191437 | 50 mm | Active Area Diameter
1982 | 1191437 | 100 µsec to 1 msec | Maximum Laser Pulse Length
1983 | 1191437 | Greater than 1J | Maximum Pulse Energy
1984 | 1191437 | 50 | Maximum Pyroelectric Rep. Rate (pulses per second)
1985 | 1191438 | 20 | Maximum Average Power (watts)
1986 | 1191438 | 2000 | Maximum Pulse Width (microseconds)
1987 | 1191438 | 12 mJ to 15J | Energy Range
1988 | 1191438 | <240 µJ | Noise Equivalent Energy
1989 | 1191438 | 0.266 to 2.1 | Wavelength Range (micrometers)
1990 | 1191438 | 35 | Maximum Beam Size (millimeters)
1991 | 1191438 | MaxBlack | Detector Coating
1992 | 1191438 | YAG | Diffuser
1993 | 1191438 | 1064 | Calibration Wavelength (nanometers)
1994 | 1191438 | ±2 | Calibration Uncertainty (%)(k=2)
1995 | 1191438 | ±3 | Energy Linearity (%)
1996 | 1191438 | 3 | Cable Length (meters)
1997 | 1191438 | 14000 (at 1064 nm, 10 ns) 2800 (at 532 nm, 10 ns) 750 (at 355 nm, 10 ns) 1000 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
1998 | 1191438 | USB | Cable Type
1999 | 1191438 | Less than 100 pps | Maximum Repetition Rate
2000 | 1191438 | 50 mm | Active Area Diameter
2001 | 1191438 | Greater than 1 msec | Maximum Laser Pulse Length
2002 | 1191438 | Greater than 1J | Maximum Pulse Energy
2003 | 1191438 | 10 | Maximum Pyroelectric Rep. Rate (pulses per second)
2004 | 1191439 | 20 | Maximum Average Power (watts)
2005 | 1191439 | 57 | Maximum Pulse Width (microseconds)
2006 | 1191439 | 2.4 mJ to 3J | Energy Range
2007 | 1191439 | <240 µJ | Noise Equivalent Energy
2008 | 1191439 | 0.266 to 2.1 | Wavelength Range (micrometers)
2009 | 1191439 | 35 | Maximum Beam Size (millimeters)
2010 | 1191439 | MaxBlack | Detector Coating
2011 | 1191439 | YAG | Diffuser
2012 | 1191439 | 1064 | Calibration Wavelength (nanometers)
2013 | 1191439 | ±2 | Calibration Uncertainty (%)(k=2)
2014 | 1191439 | ±3 | Energy Linearity (%)
2015 | 1191439 | 3 | Cable Length (meters)
2016 | 1191439 | 14000 (at 1064 nm, 10 ns) 2800 (at 532 nm, 10 ns) 750 (at 355 nm, 10 ns) 1000 (at 266 nm, 10 ns) | Maximum Energy Density (millijoules per square centimeter)
2017 | 1191439 | USB | Cable Type
2018 | 1191439 | 100 to 999 pps | Maximum Repetition Rate
2019 | 1191439 | 50 mm | Active Area Diameter
2020 | 1191439 | 10 to 99 µsec | Maximum Laser Pulse Length
2021 | 1191439 | Greater than 1J | Maximum Pulse Energy
2022 | 1191439 | 300 | Maximum Pyroelectric Rep. Rate (pulses per second)
2023 | 1191440 | 15 | Maximum Average Power (watts)
2024 | 1191440 | 1000 | Maximum Pulse Width (microseconds)
2025 | 1191440 | 3.2 mJ to 3J | Energy Range
2026 | 1191440 | <320 µJ | Noise Equivalent Energy
2027 | 1191440 | 0.5 to 3.0 | Wavelength Range (micrometers)
2028 | 1191440 | MaxBlack | Detector Coating
2029 | 1191440 | IR | Diffuser
2030 | 1191440 | 1064, 2940 | Calibration Wavelength (nanometers)
2031 | 1191440 | ±2 at 1064 nm, ±3 at 2940 nm | Calibration Uncertainty (%)(k=2)
2032 | 1191440 | ±3.5 | Energy Linearity (%)
2033 | 1191440 | 3 | Cable Length (meters)
2034 | 1191440 | >100 (at 2940 nm, 100 Î¼s) | Maximum Energy Density (millijoules per square centimeter)
2035 | 1191440 | USB | Cable Type
2036 | 1191440 | Less than 100 pps | Maximum Repetition Rate
2037 | 1191440 | 100 µsec to 1 msec | Maximum Laser Pulse Length
2038 | 1191440 | Greater than 1J | Maximum Pulse Energy
2039 | 1191440 | 500 nm to 3 µm | Wavelength Range
2040 | 1191440 | 30mm | Maximum Beam Size (millimeters)
2041 | 1191440 | 50 mm | Active Area Diameter
2042 | 1191440 | 30 | Maximum Pyroelectric Rep. Rate (pulses per second)
2043 | 1191441 | 5 | Maximum Average Power (watts)
2044 | 1191441 | 17 | Maximum Pulse Width (microseconds)
2045 | 1191441 | 50 µJ to 50 mJ | Energy Range
2046 | 1191441 | <5 µJ | Noise Equivalent Energy
2047 | 1191441 | 0.19 to 12 | Wavelength Range (micrometers)
2048 | 1191441 | MaxBlack | Detector Coating
2049 | 1191441 | No | Diffuser
2050 | 1191441 | 1064 | Calibration Wavelength (nanometers)
2051 | 1191441 | ±2 | Calibration Uncertainty (%)(k=2)
2052 | 1191441 | ±3 | Energy Linearity (%)
2053 | 1191441 | 3 | Cable Length (meters)
2054 | 1191441 | 500 | Maximum Energy Density (millijoules per square centimeter)
2055 | 1191441 | USB | Cable Type
2056 | 1191441 | 1,000 pps | Maximum Repetition Rate
2057 | 1191441 | 25 mm | Active Area Diameter
2058 | 1191441 | 10 to 99 µsec | Maximum Laser Pulse Length
2059 | 1191441 | 1 mJ to 100 mJ | Maximum Pulse Energy
2060 | 1191441 | 190 nm to 12 µm | Wavelength Range
2061 | 1191442 | 5 | Maximum Average Power (watts)
2062 | 1191442 | 17 | Maximum Pulse Width (microseconds)
2063 | 1191442 | 850 µJ to 1J | Energy Range
2064 | 1191442 | <85 µJ | Noise Equivalent Energy
2065 | 1191442 | 0.19 to 12 | Wavelength Range (micrometers)
2066 | 1191442 | MaxBlack | Detector Coating
2067 | 1191442 | No | Diffuser
2068 | 1191442 | 1064 | Calibration Wavelength (nanometers)
2069 | 1191442 | ±2 | Calibration Uncertainty (%)(k=2)
2070 | 1191442 | ±3 | Energy Linearity (%)
2071 | 1191442 | 3 | Cable Length (meters)
2072 | 1191442 | 500 | Maximum Energy Density (millijoules per square centimeter)
2073 | 1191442 | USB | Cable Type
2074 | 1191442 | 1,000 pps | Maximum Repetition Rate
2075 | 1191442 | 25 mm | Active Area Diameter
2076 | 1191442 | 10 to 99 µsec | Maximum Laser Pulse Length
2077 | 1191442 | 101 mJ to 1J | Maximum Pulse Energy
2078 | 1191442 | 190 nm to 12 µm | Wavelength Range
2079 | 1191443 | 10 | Maximum Average Power (watts)
2080 | 1191443 | 57 | Maximum Pulse Width (microseconds)
2081 | 1191443 | 400 µJ to 500 mJ | Energy Range
2082 | 1191443 | <40 µJ | Noise Equivalent Energy
2083 | 1191443 | 0.19 to 12 | Wavelength Range (micrometers)
2084 | 1191443 | MaxBlack | Detector Coating
2085 | 1191443 | No | Diffuser
2086 | 1191443 | 1064 | Calibration Wavelength (nanometers)
2087 | 1191443 | ±2 | Calibration Uncertainty (%)(k=2)
2088 | 1191443 | ±3 | Energy Linearity (%)
2089 | 1191443 | 3 | Cable Length (meters)
2090 | 1191443 | 500 | Maximum Energy Density (millijoules per square centimeter)
2091 | 1191443 | USB | Cable Type
2092 | 1191443 | 100 to 999 pps | Maximum Repetition Rate
2093 | 1191443 | 50 mm | Active Area Diameter
2094 | 1191443 | 10 to 99 µsec | Maximum Laser Pulse Length
2095 | 1191443 | 101 mJ to 1J | Maximum Pulse Energy
2096 | 1191443 | 190 nm to 12 µm | Wavelength Range
2097 | 1191444 | 10 | Maximum Average Power (watts)
2098 | 1191444 | 57 | Maximum Pulse Width (microseconds)
2099 | 1191444 | 1.6 mJ to 2J | Energy Range
2100 | 1191444 | <160 µJ | Noise Equivalent Energy
2101 | 1191444 | 0.19 to 12 | Wavelength Range (micrometers)
2102 | 1191444 | MaxBlack | Detector Coating
2103 | 1191444 | No | Diffuser
2104 | 1191444 | 1064 | Calibration Wavelength (nanometers)
2105 | 1191444 | ±2 | Calibration Uncertainty (%)(k=2)
2106 | 1191444 | ±3 | Energy Linearity (%)
2107 | 1191444 | 3 | Cable Length (meters)
2108 | 1191444 | 500 | Maximum Energy Density (millijoules per square centimeter)
2109 | 1191444 | USB | Cable Type
2110 | 1191444 | 100 to 999 pps | Maximum Repetition Rate
2111 | 1191444 | 50 mm | Active Area Diameter
2112 | 1191444 | 10 to 99 µsec | Maximum Laser Pulse Length
2113 | 1191444 | Greater than 1J | Maximum Pulse Energy
2114 | 1191444 | 190 nm to 12 µm | Wavelength Range
2115 | 1191444 | 300 | Maximum Pyroelectric Rep. Rate (pulses per second)
2116 | 1191445 | 1 | Maximum Average Power (watts)
2117 | 1191445 | 1.7 | Maximum Pulse Width (microseconds)
2118 | 1191445 | 300 nJ to 200 µJ | Energy Range
2119 | 1191445 | <30 nJ | Noise Equivalent Energy
2120 | 1191445 | Diffuse Metallic | Detector Coating
2121 | 1191445 | No | Diffuser
2122 | 1191445 | 1064 | Calibration Wavelength (nanometers)
2123 | 1191445 | ±2 | Calibration Uncertainty (%)(k=2)
2124 | 1191445 | ±3 | Energy Linearity (%)
2125 | 1191445 | 3 | Cable Length (meters)
2126 | 1191445 | 50 | Maximum Energy Density (millijoules per square centimeter)
2127 | 1191445 | USB | Cable Type
2128 | 1191445 | 10,000 pps | Maximum Repetition Rate
2129 | 1191445 | 10 mm | Active Area Diameter
2130 | 1191445 | Less than 10 µsec | Maximum Laser Pulse Length
2131 | 1191445 | 1 µJ to 1 mJ | Maximum Pulse Energy
2132 | 1191445 | 190 nm to 2100 nm | Wavelength Range
2133 | 1191445 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
2134 | 1191446 | 10 | Maximum Average Power (watts)
2135 | 1191446 | 1.7 | Maximum Pulse Width (microseconds)
2136 | 1191446 | 90 µJ to 100 mJ | Energy Range
2137 | 1191446 | <9 µJ | Noise Equivalent Energy
2138 | 1191446 | 0.19 to 2.1 | Wavelength Range (micrometers)
2139 | 1191446 | Diffuse Metallic | Detector Coating
2140 | 1191446 | No | Diffuser
2141 | 1191446 | 1064 | Calibration Wavelength (nanometers)
2142 | 1191446 | ±2 | Calibration Uncertainty (%)(k=2)
2143 | 1191446 | ±3 | Energy Linearity (%)
2144 | 1191446 | 3 | Cable Length (meters)
2145 | 1191446 | 500 | Maximum Energy Density (millijoules per square centimeter)
2146 | 1191446 | USB | Cable Type
2147 | 1191446 | 10,000 pps | Maximum Repetition Rate
2148 | 1191446 | 25 mm | Active Area Diameter
2149 | 1191446 | Less than 10 µsec | Maximum Laser Pulse Length
2150 | 1191446 | 1 mJ to 100 mJ | Maximum Pulse Energy
2151 | 1191446 | 190 nm to 2100 nm | Wavelength Range
2152 | 1191447 | 20 | Maximum Average Power (watts)
2153 | 1191447 | 1.7 | Maximum Pulse Width (microseconds)
2154 | 1191447 | 400 µJ to 1J | Minimum Energy
2155 | 1191447 | <40 µJ | Noise Equivalent Energy
2156 | 1191447 | 0.19 to 2.1 | Wavelength Range (micrometers)
2157 | 1191447 | Diffuse Metallic | Detector Coating
2158 | 1191447 | No | Diffuser
2159 | 1191447 | 1064 | Calibration Wavelength (nanometers)
2160 | 1191447 | ±2 | Calibration Uncertainty (%)(k=2)
2161 | 1191447 | ±3 | Energy Linearity (%)
2162 | 1191447 | 3 | Cable Length (meters)
2163 | 1191447 | 500 | Maximum Energy Density (millijoules per square centimeter)
2164 | 1191447 | USB | Cable Type
2165 | 1191447 | 10,000 pps | Maximum Repetition Rate
2166 | 1191447 | 50 mm | Active Area Diameter
2167 | 1191447 | Less than 10 µsec | Maximum Laser Pulse Length
2168 | 1191447 | 101 mJ to 1J | Maximum Pulse Energy
2169 | 1191447 | 190 nm to 2100 nm | Wavelength Range
2170 | 1191447 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
2171 | 1191448 | 5 | Maximum Average Power (watts)
2172 | 1191448 | 43 | Maximum Pulse Width (microseconds)
2173 | 1191448 | 90 µJ to 100 mJ | Energy Range
2174 | 1191448 | <9 µJ | Noise Equivalent Energy
2175 | 1191448 | 0.19 to 2.1 | Wavelength Range (micrometers)
2176 | 1191448 | 200 | Maximum Energy Density (millijoules per square centimeter)
2177 | 1191448 | MaxUV | Detector Coating
2178 | 1191448 | No | Diffuser
2179 | 1191448 | 193 | Calibration Wavelength (nanometers)
2180 | 1191448 | ±3 | Calibration Uncertainty (%)(k=2)
2181 | 1191448 | ±3 | Energy Linearity (%)
2182 | 1191448 | 3 | Cable Length (meters)
2183 | 1191448 | USB | Cable Type
2184 | 1191448 | 100 to 999 pps | Maximum Repetition Rate
2185 | 1191448 | 25 mm | Active Area Diameter
2186 | 1191448 | 10 to 99 µsec | Maximum Laser Pulse Length
2187 | 1191448 | 1 mJ to 100 mJ | Maximum Pulse Energy
2188 | 1191448 | 190 nm to 2100 nm | Wavelength Range
2189 | 1191448 | 500 | Maximum Pyroelectric Rep. Rate (pulses per second)
2190 | 1191449 | 15 | Maximum Average Power (watts)
2191 | 1191449 | 86 | Maximum Pulse Width (microseconds)
2192 | 1191449 | 800 µJ to 1J | Energy Range
2193 | 1191449 | <80 µJ | Noise Equivalent Energy
2194 | 1191449 | 0.19 to 0.266 | Wavelength Range (micrometers)
2195 | 1191449 | MaxUV | Detector Coating
2196 | 1191449 | Yes | Diffuser
2197 | 1191449 | 193 | Calibration Wavelength (nanometers)
2198 | 1191449 | ±3 | Calibration Uncertainty (%)(k=2)
2199 | 1191449 | ±3 | Energy Linearity (%)
2200 | 1191449 | 3 | Cable Length (meters)
2201 | 1191449 | 200 | Maximum Energy Density (millijoules per square centimeter)
2202 | 1191449 | USB | Cable Type
2203 | 1191449 | 100 to 999 pps | Maximum Repetition Rate
2204 | 1191449 | 50 mm | Active Area Diameter
2205 | 1191449 | 10 to 99 µsec | Maximum Laser Pulse Length
2206 | 1191449 | 100 µsec to 1 msec | Maximum Laser Pulse Length
2207 | 1191449 | 101 mJ to 1J | Maximum Pulse Energy
2208 | 1191449 | 190 nm to 266 nm | Wavelength Range
2209 | 1214871 | RS232 | Cable Type
2210 | 1214871 | 150 | Minimum Wavelength (nanometers)
2211 | 1214871 | 1000 | Maximum Wavelength (nanometers)
2212 | 1214871 | 50 | Minimum Power (watts)
2213 | 1214871 | 1000 | Maximum Power (watts)
2214 | 1214871 | Water | Cooling Type
2215 | 1214871 | 0.15 to 1 µm | Wavelength Range (micrometers)
2216 | 1214871 | 50 to 1000W | Power Range
2217 | 1214871 | 1 to 2.5 kW/cm² - power dependent | Maximum Power Density (kilowatts per square centimeter)
2218 | 1214871 | 600 mJ/cm² | Maximum Energy Density (millijoules per square centimeter)
2219 | 1214871 | ±3 | Calibration Uncertainty (%)(k=2)
2220 | 1214871 | ±1 | Power Linearity (%)
2221 | 1214871 | ±1.5 | Spectral Compensation Accuracy (%)
2222 | 1214871 | 1070 | Calibration Wavelength (nanometers)
2223 | 1214871 | Water | Cooling Method
2224 | 1214871 | 0.3 | Cable Length (meters)
2225 | 1214871 | None | Window
2226 | 1214871 | None | Diffuser
2227 | 1214871 | 100 mm | Active Area Diameter
2228 | 1214871 | 20 | Noise Equivalent Power (milliwatts)
2229 | 1214871 | 4 | Response Time, Speed-up On, 0% to 95% (seconds)
2230 | 1214871 | UV | Detector Coating
2231 | 1219962 | 20 | Maximum Average Power (watts)
2232 | 1219962 | 2000 | Maximum Pulse Width (microseconds)
2233 | 1219962 | 12 mJ to 15J | Energy Range
2234 | 1219962 | <240 µJ | Noise Equivalent Energy
2235 | 1219962 | 0.266 to 2.1 | Wavelength Range (micrometers)
2236 | 1219962 | MaxBlack | Detector Coating
2237 | 1219962 | YAG | Diffuser
2238 | 1219962 | 1064 | Calibration Wavelength (nanometers)
2239 | 1219962 | ±2 | Calibration Uncertainty (%)(k=2)
2240 | 1219962 | ±3 | Energy Linearity (%)
2241 | 1219962 | 3 | Cable Length (meters)
2242 | 1219962 | 14.0 (at 1064 nm, 10 ns); 2.8 (at 532 nm); 0.75 (at 355 nm); 1.0 (at 266 nm) | Maximum Energy Density (millijoules per square centimeter)
2243 | 1219962 | RS232 | Cable Type
2244 | 1219962 | Less than 100 pps | Maximum Repetition Rate
2245 | 1219962 | Greater than 1 msec | Maximum Laser Pulse Length
2246 | 1219962 | Greater than 1J | Maximum Pulse Energy
2247 | 1219962 | 35 | Maximum Beam Size (millimeters)
2248 | 1219962 | 50 mm | Active Area Diameter
2249 | 1219962 | 50 | Maximum Pyroelectric Rep. Rate (pulses per second)
2250 | 1223336 | 600 | Maximum Energy Density (millijoules per square centimeter)
2251 | 1223336 | USB | Cable Type
2252 | 1223336 | 190 | Minimum Wavelength (nanometers)
2253 | 1223336 | 11000 | Maximum Wavelength (nanometers)
2254 | 1223336 | 0.3 | Minimum Power (watts)
2255 | 1223336 | 150 | Maximum Power (watts)
2256 | 1223336 | 150 | Maximum Long-Pulse Energy (joules)
2257 | 1223336 | Air | Cooling Type
2258 | 1223336 | 0.19 to 11 µm | Wavelength Range (micrometers)
2259 | 1223336 | 300 mW to 150W | Power Range
2260 | 1223336 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
2261 | 1223336 | 1 to 150 | Long-Pulse Joules Range (joules)
2262 | 1223336 | 1 | Noise Equivalent Power (milliwatts)
2263 | 1223336 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2264 | 1223336 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2265 | 1223336 | Broadband | Detector Coating
2266 | 1223336 | ±2 | Calibration Uncertainty (%)(k=2)
2267 | 1223336 | ±1 | Power Linearity (%)
2268 | 1223336 | ±1.5 | Spectral Compensation Accuracy (%)
2269 | 1223336 | 10600 | Calibration Wavelength (nanometers)
2270 | 1223336 | Air | Cooling Method
2271 | 1223336 | 2.5 | Cable Length (meters)
2272 | 1223336 | None | Window
2273 | 1223336 | None | Diffuser
2274 | 1223336 | 50 mm | Active Area Diameter
2275 | 1224012 | 10µW to 1W | Power Range
2277 | 1224012 | 190 | Minimum Wavelength (nanometers)
2278 | 1224012 | 1100 | Maximum Wavelength (nanometers)
2279 | 1224012 | 3 | Minimum Beam Size (micrometers)
2280 | 1224012 | .19 to 1.1 | Wavelength Range (micrometers)
2281 | 1224012 | CW Lasers Only | Capture Modes
2282 | 1224012 | 5 mm circular | Sensor Aperture
2283 | 1224012 | 190 nm to 1100 nm | Wavelength Range
2284 | 1224012 | Knife-edge (CW Only) | Profiler Type
2285 | 1224012 | 1 µm for beams >100 µm in size (0.1µm for beams <100 µm in size) | Beam Size Resolution
2286 | 1224012 | 1 | Position Measurement Resolution (micrometers)
2287 | 1224012 | ±15 | Position Measurement Accuracy (micrometers)
2288 | 1224012 | ±2 | Beam Width Measurement Accuracy (%)
2289 | 1224012 | 5 mm | Active Area Diameter
2290 | 1224012 | 5 mm (diameter) | Active Area Size
2291 | 1224012 | 3 microns to 4 mm | Recommended Beam Size
2292 | 1224014 | 10 µW to 1W | Power Range
2294 | 1224014 | 190 | Minimum Wavelength (nanometers)
2295 | 1224014 | 1100 | Maximum Wavelength (nanometers)
2296 | 1224014 | 15 | Minimum Beam Size (micrometers)
2297 | 1224014 | .19 to 1.1 | Wavelength Range (micrometers)
2298 | 1224014 | CW Lasers Only | Capture Modes
2299 | 1224014 | 9 mm square | Sensor Aperture
2300 | 1224014 | 1 µm for beams >100 µm in size (0.1 µm for beams <100 µm in size) | Beam Size Resolution
2301 | 1224014 | 1 | Position Measurement Resolution (micrometers)
2302 | 1224014 | ±15 | Position Measurement Accuracy (micrometers)
2303 | 1224014 | ±2 | Beam Width Measurement Accuracy (%)
2304 | 1224014 | 190 nm to 1100 nm | Wavelength Range
2305 | 1224014 | Knife-edge (CW Only) | Profiler Type
2306 | 1224014 | 9 mm | Active Area Diameter
2307 | 1224014 | 9 x 9 mm | Active Area Size
2308 | 1224014 | 15 microns to 8 mm | Recommended Beam Size
2309 | 1224016 | 10µW to 1W | Power Range
2311 | 1224016 | 800 nm to 1800 nm | Wavelength Range
2312 | 1224016 | 15 microns to 2 mm | Recommended Beam Size
2313 | 1224016 | Knife-edge (CW Only) | Profiler Type
2314 | 1224016 | 800 | Minimum Wavelength (nanometers)
2315 | 1224016 | 1800 | Maximum Wavelength (nanometers)
2316 | 1224016 | 15 | Minimum Beam Size (micrometers)
2317 | 1224016 | .8 to 1.8 | Wavelength Range (micrometers)
2318 | 1224016 | CW Lasers Only | Capture Modes
2319 | 1224016 | 3 mm circular | Sensor Aperture
2320 | 1224016 | 1 µm for beams >100 µm in size (0.1 µm for beams <100 µm in size) | Beam Size Resolution
2321 | 1224016 | 1 | Position Measurement Resolution (micrometers)
2322 | 1224016 | ±15 | Position Measurement Accuracy (micrometers)
2323 | 1224016 | ±2 | Beam Width Measurement Accuracy (%)
2324 | 1224016 | 3 mm | Active Area Diameter
2325 | 1224016 | 3 mm (diameter) | Active Area Size
2326 | 1224018 | 10µW to 1W | Power Range
2328 | 1224018 | 800 nm to 1800 nm | Wavelength Range
2329 | 1224018 | 15 microns to 2 mm | Recommended Beam Size
2330 | 1224018 | Knife-edge (CW Only) | Profiler Type
2331 | 1224018 | 800 | Minimum Wavelength (nanometers)
2332 | 1224018 | 1800 | Maximum Wavelength (nanometers)
2333 | 1224018 | 15 | Minimum Beam Size (micrometers)
2334 | 1224018 | .8 to 1.8 | Wavelength Range (micrometers)
2335 | 1224018 | CW Lasers Only | Capture Modes
2336 | 1224018 | 3 mm circular | Sensor Aperture
2337 | 1224018 | 1 µm for beams >100 µm in size (0.1 µm for beams <100 µm in size) | Beam Size Resolution
2338 | 1224018 | 1 | Position Measurement Resolution (micrometers)
2339 | 1224018 | ±15 | Position Measurement Accuracy (micrometers)
2340 | 1224018 | ±2 | Beam Width Measurement Accuracy (%)
2341 | 1224018 | 3 mm | Active Area Diameter
2342 | 1224018 | 3 mm (diameter) | Active Area Size
2343 | 1224020 | 10 µW to 1W | Power Range
2345 | 1224020 | 800 nm to 1800 nm | Wavelength Range
2346 | 1224020 | 3 microns to 2 mm | Recommended Beam Size
2347 | 1224020 | Knife-edge (CW Only) | Profiler Type
2348 | 1224020 | 800 | Minimum Wavelength (nanometers)
2349 | 1224020 | 1800 | Maximum Wavelength (nanometers)
2350 | 1224020 | 3 | Minimum Beam Size (micrometers)
2351 | 1224020 | .8 to 1.8 | Wavelength Range (micrometers)
2352 | 1224020 | CW Lasers Only | Capture Modes
2353 | 1224020 | 3 mm circular | Sensor Aperture
2354 | 1224020 | 1 µm for beams >100 µm in size (0.1 µm for beams <100 µm in size) | Beam Size Resolution
2355 | 1224020 | 1 | Position Measurement Resolution (micrometers)
2356 | 1224020 | ±15 | Position Measurement Accuracy (micrometers)
2357 | 1224020 | ±2 | Beam Width Measurement Accuracy (%)
2358 | 1224020 | 3 mm | Active Area Diameter
2359 | 1224020 | 3 mm (diameter) | Active Area Size
2360 | 1230323 | 600 | Maximum Energy Density (millijoules per square centimeter)
2361 | 1230323 | RS232 | Cable Type
2362 | 1230323 | 150 | Minimum Wavelength (nanometers)
2363 | 1230323 | 1000 | Maximum Wavelength (nanometers)
2364 | 1230323 | 0.005 | Minimum Power (watts)
2365 | 1230323 | 2 | Maximum Power (watts)
2366 | 1230323 | 2 | Maximum Long-Pulse Energy (joules)
2367 | 1230323 | Air | Cooling Type
2368 | 1230323 | 0.15 to 1 µm | Wavelength Range (micrometers)
2369 | 1230323 | 5 mW to 2W | Power Range
2370 | 1230323 | 5 | Maximum Intermittent Power (watts)(<5 minutes)
2371 | 1230323 | 0.5 to 2 | Long-Pulse Joules Range (joules)
2372 | 1230323 | 0.2 | Noise Equivalent Power (milliwatts)
2373 | 1230323 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2374 | 1230323 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
2375 | 1230323 | UV | Detector Coating
2376 | 1230323 | ±1 | Calibration Uncertainty (%)(k=2)
2377 | 1230323 | ±1 | Power Linearity (%)
2378 | 1230323 | ±1.5 | Spectral Compensation Accuracy (%)
2379 | 1230323 | 514 | Calibration Wavelength (nanometers)
2380 | 1230323 | Air | Cooling Method
2381 | 1230323 | 0.3 | Cable Length (meters)
2382 | 1230323 | None | Window
2383 | 1230323 | None | Diffuser
2384 | 1230323 | 19 mm | Active Area Diameter
2385 | 1235755 | USB | Cable Type
2386 | 1235755 | 250 | Minimum Wavelength (nanometers)
2387 | 1235755 | 10600 | Maximum Wavelength (nanometers)
2388 | 1235755 | 50 | Minimum Power (watts)
2389 | 1235755 | 5000 | Maximum Power (watts)
2390 | 1235755 | Water | Cooling Type
2391 | 1235755 | 0.25 to 10.6 µm | Wavelength Range (micrometers)
2392 | 1235755 | 100 to 5000W | Power Range
2393 | 1235755 | 1 to 2.5 kW/cm² - power dependent | Maximum Power Density (kilowatts per square centimeter)
2394 | 1235755 | 600 | Maximum Energy Density (millijoules per square centimeter)
2395 | 1235755 | ±5 | Calibration Uncertainty (%)(k=2)
2396 | 1235755 | ±1 | Power Linearity (%)
2397 | 1235755 | ±1.5 | Spectral Compensation Accuracy (%)
2398 | 1235755 | 1070 | Calibration Wavelength (nanometers)
2399 | 1235755 | Water | Cooling Method
2400 | 1235755 | 2.5 | Cable Length (meters)
2401 | 1235755 | None | Window
2402 | 1235755 | None | Diffuser
2403 | 1235755 | 100 mm | Active Area Diameter
2404 | 1235755 | 20 | Noise Equivalent Power (milliwatts)
2405 | 1235755 | 4 | Response Time, Speed-up On, 0% to 95% (seconds)
2406 | 1235755 | Broadband | Detector Coating
2407 | 1236447 | 600 | Maximum Energy Density (millijoules per square centimeter)
2408 | 1236447 | RS232 | Cable Type
2409 | 1236447 | 150 | Minimum Wavelength (nanometers)
2410 | 1236447 | 1000 | Maximum Wavelength (nanometers)
2411 | 1236447 | 0.005 | Minimum Power (watts)
2412 | 1236447 | 10 | Maximum Power (watts)
2413 | 1236447 | 10 | Maximum Long-Pulse Energy (joules)
2414 | 1236447 | Air | Cooling Type
2415 | 1236447 | 0.15 to 1 µm | Wavelength Range (micrometers)
2416 | 1236447 | 5 mW to 10W | Power Range
2417 | 1236447 | 30 | Maximum Intermittent Power (watts)(<5 minutes)
2418 | 1236447 | 0.5 to 10 | Long-Pulse Joules Range (joules)
2419 | 1236447 | 0.2 | Noise Equivalent Power (milliwatts)
2420 | 1236447 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2421 | 1236447 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
2422 | 1236447 | UV | Detector Coating
2423 | 1236447 | ±1 | Calibration Uncertainty (%)(k=2)
2424 | 1236447 | ±1 | Power Linearity (%)
2425 | 1236447 | ±1.5 | Spectral Compensation Accuracy (%)
2426 | 1236447 | 514 | Calibration Wavelength (nanometers)
2427 | 1236447 | Air | Cooling Method
2428 | 1236447 | 0.3 | Cable Length (meters)
2429 | 1236447 | None | Window
2430 | 1236447 | None | Diffuser
2431 | 1236447 | 19 mm | Active Area Diameter
2432 | 1257617 | 600 | Maximum Energy Density (millijoules per square centimeter)
2433 | 1257617 | USB | Cable Type
2434 | 1257617 | 150 | Minimum Wavelength (nanometers)
2435 | 1257617 | 1000 | Maximum Wavelength (nanometers)
2436 | 1257617 | 0.005 | Minimum Power (watts)
2437 | 1257617 | 2 | Maximum Power (watts)
2438 | 1257617 | 2 | Maximum Long-Pulse Energy (joules)
2439 | 1257617 | Air | Cooling Type
2440 | 1257617 | 0.15 to 1 µm | Wavelength Range (micrometers)
2441 | 1257617 | 5 mW to 2W | Power Range
2442 | 1257617 | 5 | Maximum Intermittent Power (watts)(<5 minutes)
2443 | 1257617 | 0.5 to 2 | Long-Pulse Joules Range (joules)
2444 | 1257617 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2445 | 1257617 | UV | Detector Coating
2446 | 1257617 | ±1 | Calibration Uncertainty (%)(k=2)
2447 | 1257617 | ±1 | Power Linearity (%)
2448 | 1257617 | ±1.5 | Spectral Compensation Accuracy (%)
2449 | 1257617 | 514 | Calibration Wavelength (nanometers)
2450 | 1257617 | Air | Cooling Method
2451 | 1257617 | 2.5 | Cable Length (meters)
2452 | 1257617 | None | Window
2453 | 1257617 | None | Diffuser
2454 | 1257617 | 19 mm | Active Area Diameter
2455 | 1257617 | 0.2 | Noise Equivalent Power (milliwatts)
2456 | 1257617 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
2457 | 1263294 | 600 | Maximum Energy Density (millijoules per square centimeter)
2458 | 1263294 | USB | Cable Type
2459 | 1263294 | 150 | Minimum Wavelength (nanometers)
2460 | 1263294 | 1000 | Maximum Wavelength (nanometers)
2461 | 1263294 | 0.01 | Minimum Power (watts)
2462 | 1263294 | 30 | Maximum Power (watts)
2463 | 1263294 | 50 | Maximum Long-Pulse Energy (joules)
2464 | 1263294 | Air | Cooling Type
2465 | 1263294 | 0.15 to 1 µm | Wavelength Range (micrometers)
2466 | 1263294 | 10 mW to 30W | Power Range
2467 | 1263294 | 50 | Maximum Intermittent Power (watts)(<5 minutes)
2468 | 1263294 | 0.5 to 50 | Long-Pulse Joules Range (joules)
2469 | 1263294 | 0.5 | Noise Equivalent Power (milliwatts)
2470 | 1263294 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2471 | 1263294 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
2472 | 1263294 | UV | Detector Coating
2473 | 1263294 | ±1 | Calibration Uncertainty (%)(k=2)
2474 | 1263294 | ±1 | Power Linearity (%)
2475 | 1263294 | ±1.5 | Spectral Compensation Accuracy (%)
2476 | 1263294 | 514 | Calibration Wavelength (nanometers)
2477 | 1263294 | Air | Cooling Method
2478 | 1263294 | 2.5 | Cable Length (meters)
2479 | 1263294 | None | Window
2480 | 1263294 | None | Diffuser
2481 | 1263294 | 19 mm | Active Area Diameter
2482 | 1268881 | Trending, tuning, histogram, data logging, statistics | Statistical Analysis
2483 | 1268881 | 0 to 1, 2, or 4 (selectable) | Analog Output Text
2484 | 1268881 | 10 Hz, 20 kHz (High Speed), 625 kHz (Snapshot) | Power Sampling Rate
2485 | 1268881 | 12VDC - center positive; (90 to 260 VAC, 50/60 Hz AC power adapter included) | Instrument Power
2486 | 1268881 | CE, RoHS, WEEE, ISO 17025 | Compliance
2488 | 1268881 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
2489 | 1268881 | Optical | Sensor Compatibility
2490 | 1268881 | PowerMax-Pro | Sensor Compatibility
2491 | 1268881 | Pyroelectric | Sensor Compatibility
2492 | 1268881 | Thermopile (LM-model with beam position) | Sensor Compatibility
2493 | 1268881 | Thermopile (PM-model) | Sensor Compatibility
2494 | 1268881 | Energy (pulse-to-pulse) | Measurement Type
2495 | 1268881 | Energy (single-shot) | Measurement Type
2496 | 1268881 | Power | Measurement Type
2497 | 1268881 | Position | Measurement Type
2498 | 1268881 | Analog output | Outputs & Interfaces
2499 | 1268881 | USB | Outputs & Interfaces
2500 | 1268881 | RS232 | Outputs & Interfaces
2501 | 1268881 | Yes | Thermopile sensor compatibility
2502 | 1268881 | Yes | Optical sensor compatibility
2503 | 1268881 | Yes | Pyroelectric sensor compatibility
2504 | 1268881 | Yes | PowerMax-Pro sensor compatibility
2505 | 1268881 | Yes | USB Interface
2506 | 1268881 | Yes | RS232 Interface
2507 | 1268881 | Yes | Analog output
2508 | 1268881 | Yes | External Trigger Input
2509 | 1268881 | None | Display Type
2510 | 1268881 | No | Battery Operation
2511 | 1268881 | Sensor dependent (reference sensor specifications) | Measurement Range
2512 | 1268881 | 0.1% of full scale | Measurement Resolution
2513 | 1268881 | 20,000 Hz every pulse | Maximum Pulse Energy Repetition Rate (hertz)
2514 | 1268881 | Meter accuracy + sensor accuracy | Accuracy - System
2515 | 1268881 | ±1.0% | Accuracy - Analog Output
2516 | 1268881 | ±1.0% | Calibration Uncertainty (%)(k=2)
2517 | 1268881 | PC Only | Display
2518 | 1268881 | USB with Virtual COM Port, RS-232, LabVIEW, PC application | Computer Interface
2519 | 1268881 | External and internal trigger (selectable) | Trigger
2520 | 1268881 | 105 x 105 x 32 mm (4.1 x 4.1 x 1.3 in.) | Size (H x W x D)
2521 | 1268881 | 0.3 kg (0.6 lbs.) | Weight
2522 | 1268881 | Yes | C24
2523 | 1282868 | Continuous (CW), pulsed | Capture Modes
2524 | 1282868 | 1 msec to 500 msec, default at 5 msec | Variable Exposure Time
2526 | 1282868 | 400 | Minimum Wavelength (nanometers)
2527 | 1282868 | 1100 | Maximum Wavelength (nanometers)
2528 | 1282868 | 0.15 to 4.0 | Recommended Beam Diameters (millimeters)
2529 | 1282868 | 12-bit | Camera Bit Depth
2530 | 1282868 | 1280 x 1024 | Sensor Elements (pixels)
2531 | 1282868 | n/a | Effective Pixel Resolution (micrometers)
2532 | 1282868 | 4.6 x 4.6 | Pixel Size (micrometers)
2533 | 1282868 | 5.9 x 4.8 | Sensor Active Area (millimeters)(H x V)
2534 | 1282868 | 190 to 1100 | Spectral Range (nanometers) without LDFP
2535 | 1282868 | 400 to 1100 | Spectral Range (nanometers) with LDFP included
2536 | 1282868 | 190 to 355 | Spectral Range (nanometers) with BIP-12F accessory
2537 | 1282868 | 75 | Trigger Delay (microseconds)
2538 | 1282868 | 200 | Maximum Pulse Trigger in Rate (hertz)
2539 | 1282868 | 32 mJ/cm2 at 1064 nm | Damage Threshold without LDFP
2540 | 1282868 | 13 mW/cm2 at 633 nm | CW Saturation with LDFP at 633 nanometers
2541 | 1282868 | 0.005 mW/cm2 at 633 nm | CW Saturation without LDFP at 633 nanometers
2542 | 1282868 | 70 mW/cm2 at 1064 nm | CW Saturation with LDFP at 1064 nanometers
2543 | 1282868 | 0.34 mW/cm2 at 1064 nm | CW Saturation without LDFP at 1064 nanometers
2544 | 1282868 | 0.4 mJ/cm2 at 1064 nm | Pulsed Saturation with LDFP
2545 | 1282868 | .002 mJ/cm2 at 1064 nm | Pulsed Saturation without LDFP
2546 | 1282868 | Connector BNC receptacle (trigger cable included) | Trigger
2547 | 1282868 | 400 nm to 1100 nm | Wavelength Range
2548 | 1282868 | 150 microns to 4 mm | Recommended Beam Size
2549 | 1282868 | Camera (CW or Pulsed) | Profiler Type
2550 | 1282868 | 5.9 x 4.8 mm (1/2 inch format) | Active Area Size
2551 | 1282868 | LCAM20 | Coupon Spec
2552 | 1282868 | Yes | C24
2553 | 1282870 | Continuous (CW), pulsed | Capture Modes
2554 | 1282870 | 1 msec to 500 msec, default at 5 msec | Variable Exposure Time
2556 | 1282870 | 400 | Minimum Wavelength (nanometers)
2557 | 1282870 | 1100 | Maximum Wavelength (nanometers)
2558 | 1282870 | 0.2 to 6.0 | Recommended Beam Diameters (millimeters)
2559 | 1282870 | 14-bit | Camera Bit Depth
2560 | 1282870 | 1280 x 1024 | Sensor Elements (pixels)
2561 | 1282870 | n/a | Effective Pixel Resolution (micrometers)
2562 | 1282870 | 6.5 x 6.5 | Pixel Size (micrometers)
2563 | 1282870 | 8.3 x 6.6 | Sensor Active Area (millimeters)(H x V)
2564 | 1282870 | 190 to 1100 | Spectral Range (nanometers) without LDFP
2565 | 1282870 | 400 to 1100 | Spectral Range (nanometers) with LDFP included
2566 | 1282870 | 190 to 355 | Spectral Range (nanometers) with BIP-12F accessory
2567 | 1282870 | 20 | Trigger Delay (microseconds)
2568 | 1282870 | 200 | Maximum Pulse Trigger in Rate (hertz)
2569 | 1282870 | 32 mJ/cm2 at 1064 nm | Damage Threshold without LDFP
2570 | 1282870 | 5 mW/cm2 at 633 nm | CW Saturation with LDFP at 633 nanometers
2571 | 1282870 | .002 mW/cm2 at 633 nm | CW Saturation without LDFP at 633 nanometers
2572 | 1282870 | 25 mW/cm2 at 1064 nm | CW Saturation with LDFP at 1064 nanometers
2573 | 1282870 | .125 mW/cm2 at 1064 nm | CW Saturation without LDFP at 1064 nanometers
2574 | 1282870 | 0.15 mJ/cm2 at 1064 nm | Pulsed Saturation with LDFP
2575 | 1282870 | 0.007 mJ/cm2 at 1064 nm | Pulsed Saturation without LDFP
2576 | 1282870 | Connector BNC receptacle (trigger cable included) | Trigger
2577 | 1282870 | 400 nm to 1100 nm | Wavelength Range
2578 | 1282870 | 200 microns to 6 mm | Recommended Beam Size
2579 | 1282870 | Camera (CW or Pulsed) | Profiler Type
2580 | 1282870 | 8.3 x 6.6 mm (2/3 inch format) | Active Area Size
2581 | 1282870 | LCAM20 | Coupon Spec
2582 | 1282870 | Yes | C24
2583 | 1286950 | 0.06 | Maximum Average Power (watts)
2584 | 1286950 | 1 | Maximum Pulse Width (microseconds)
2585 | 1286950 | 1 nJ to 600 nJ (at 1064 nm) | Energy Range
2586 | 1286950 | <100 (at 1064 nm) | Noise Equivalent Energy
2587 | 1286950 | .800 to 1.700 | Wavelength Range (micrometers)
2588 | 1286950 | ND2 | Diffuser
2589 | 1286950 | 1064 | Calibration Wavelength (nanometers)
2590 | 1286950 | ±3 | Calibration Uncertainty (%)(k=2)
2591 | 1286950 | ±3 | Energy Linearity (%)
2592 | 1286950 | 3 | Cable Length (meters)
2593 | 1286950 | USB | Cable Type
2594 | 1286950 | 10,000 pps | Maximum Repetition Rate
2595 | 1286950 | 10 mm | Active Area Diameter
2596 | 1286950 | Less than 10 µsec | Maximum Laser Pulse Length
2597 | 1286950 | Less than 1 µJ | Maximum Pulse Energy
2598 | 1286950 | 800 nm to 1700 nm | Wavelength Range
2599 | 1286950 | Germanium | Detector
2600 | 1286950 | 10000 | Maximum Pyroelectric Rep. Rate (pulses per second)
2601 | 1287077 | 50 | Maximum Energy Density (millijoules per square centimeter)
2602 | 1287077 | USB | Cable Type
2603 | 1287077 | 300 | Minimum Wavelength (nanometers)
2604 | 1287077 | 2100 | Maximum Wavelength (nanometers)
2605 | 1287077 | 0.0001 | Minimum Power (watts)
2606 | 1287077 | 1 | Maximum Power (watts)
2607 | 1287077 | 1 | Maximum Long-Pulse Energy (joules)
2608 | 1287077 | Air | Cooling Type
2609 | 1287077 | 0.3 to 2.1 µm | Wavelength Range (micrometers)
2610 | 1287077 | 100 µW to 1W | Power Range
2611 | 1287077 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
2612 | 1287077 | 0.001 to 1 | Long-Pulse Joules Range (joules)
2613 | 1287077 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
2614 | 1287077 | ±1 | Calibration Uncertainty (%)(k=2)
2615 | 1287077 | ±1 | Power Linearity (%)
2616 | 1287077 | ±1.5 | Spectral Compensation Accuracy (%)
2617 | 1287077 | 514 | Calibration Wavelength (nanometers)
2618 | 1287077 | Air | Cooling Method
2619 | 1287077 | 2.5 | Cable Length (meters)
2620 | 1287077 | Yes | Window
2621 | 1287077 | None | Diffuser
2622 | 1287077 | 10 mm | Active Area Diameter
2623 | 1287077 | 0.003 | Noise Equivalent Power (milliwatts)
2624 | 1287077 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
2625 | 1287077 | Black | Detector Coating
2626 | 1288940 | Water, Air (intermittent) | Cooling Type
2627 | 1288940 | 100, <5 min. | Maximum Intermittent Power (watts)(<5 minutes)
2628 | 1288940 | 600 | Maximum Energy Density (millijoules per square centimeter)
2629 | 1288940 | Water, Air (intermittent) | Cooling Method
2630 | 1288940 | USB | Cable Type
2631 | 1288940 | 190 | Minimum Wavelength (nanometers)
2632 | 1288940 | 11000 | Maximum Wavelength (nanometers)
2633 | 1288940 | 0.3 | Minimum Power (watts)
2634 | 1288940 | 100 | Maximum Power (watts)
2635 | 1288940 | 150 | Maximum Long-Pulse Energy (joules)
2636 | 1288940 | 0.19 to 11 µm | Wavelength Range (micrometers)
2637 | 1288940 | 300 mW to 100W (water-cooled) | Power Range
2638 | 1288940 | 1 to 100 | Long-Pulse Joules Range (joules)
2639 | 1288940 | 1 | Noise Equivalent Power (milliwatts)
2640 | 1288940 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2641 | 1288940 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
2642 | 1288940 | Broadband | Detector Coating
2643 | 1288940 | ±2 | Calibration Uncertainty (%)(k=2)
2644 | 1288940 | ±1 | Power Linearity (%)
2645 | 1288940 | ±1.5 | Spectral Compensation Accuracy (%)
2646 | 1288940 | 10600 | Calibration Wavelength (nanometers)
2647 | 1288940 | 2.5 | Cable Length (meters)
2648 | 1288940 | None | Window
2649 | 1288940 | None | Diffuser
2650 | 1288940 | 19 mm | Active Area Diameter
2651 | 1288992 | 50 | Maximum Energy Density (millijoules per square centimeter)
2652 | 1288992 | RS232 | Cable Type
2653 | 1288992 | 300 | Minimum Wavelength (nanometers)
2654 | 1288992 | 2100 | Maximum Wavelength (nanometers)
2655 | 1288992 | 0.0001 | Minimum Power (watts)
2656 | 1288992 | 1 | Maximum Power (watts)
2657 | 1288992 | 1 | Maximum Long-Pulse Energy (joules)
2658 | 1288992 | Air | Cooling Type
2659 | 1288992 | 0.3 to 2.1 µm | Wavelength Range (micrometers)
2660 | 1288992 | 100 µW to 1W | Power Range
2661 | 1288992 | 3 | Maximum Intermittent Power (watts)(<5 minutes)
2662 | 1288992 | 0.001 to 1 | Long-Pulse Joules Range (joules)
2663 | 1288992 | 500 W/cm² | Maximum Power Density (kilowatts per square centimeter)
2664 | 1288992 | ±1 | Calibration Uncertainty (%)(k=2)
2665 | 1288992 | ±1 | Power Linearity (%)
2666 | 1288992 | ±1.5 | Spectral Compensation Accuracy (%)
2667 | 1288992 | 514 | Calibration Wavelength (nanometers)
2668 | 1288992 | Air | Cooling Method
2669 | 1288992 | 0.3 | Cable Length (meters)
2670 | 1288992 | Yes | Window
2671 | 1288992 | None | Diffuser
2672 | 1288992 | 10 mm | Active Area Diameter
2673 | 1288992 | 0.003 | Noise Equivalent Power (milliwatts)
2674 | 1288992 | 2 | Response Time, Speed-up On, 0% to 95% (seconds)
2675 | 1288992 | Black | Detector Coating
2676 | 1289935 | 10 | Maximum Average Power (watts)
2677 | 1289935 | 86 | Maximum Pulse Width (microseconds)
2678 | 1289935 | 200 µJ to 250 mJ | Energy Range
2679 | 1289935 | <20 µJ | Noise Equivalent Energy
2680 | 1289935 | 0.19 to 2.1 | Wavelength Range (micrometers)
2681 | 1289935 | 200 | Maximum Energy Density (millijoules per square centimeter)
2682 | 1289935 | MaxUV | Detector Coating
2683 | 1289935 | No | Diffuser
2684 | 1289935 | 193 | Calibration Wavelength (nanometers)
2685 | 1289935 | ±3 | Calibration Uncertainty (%)(k=2)
2686 | 1289935 | ±3 | Energy Linearity (%)
2687 | 1289935 | 3 | Cable Length (meters)
2688 | 1289935 | USB | Cable Type
2689 | 1289935 | 100 to 999 pps | Maximum Repetition Rate
2690 | 1289935 | 50 mm | Active Area Diameter
2691 | 1289935 | 10 to 99 µsec | Maximum Laser Pulse Length
2692 | 1289935 | 100 µsec to 1 msec | Maximum Laser Pulse Length
2693 | 1289935 | 101 mJ to 1J | Maximum Pulse Energy
2694 | 1289935 | 190 nm to 2100 nm | Wavelength Range
2695 | 1289935 | 200 | Maximum Pyroelectric Rep. Rate (pulses per second)
2696 | 1299161 | 325 | Minimum Wavelength (nanometers)
2697 | 1299161 | 1065 | Maximum Wavelength (nanometers)
2698 | 1299161 | 85000 | Minimum Power (nanowatts)
2699 | 1299161 | 140 | Max Power (milliwatts)
2700 | 1299161 | 325 nm to 1065 nm | Wavelength Range (micrometers)
2701 | 1299161 | 8.5 µW to > 140 mW | Power Range
2702 | 1299161 | 170 | Noise Equivalent Power (nanowatts)
2703 | 1299161 | 20 | Maximum Power Density (watts per square centimeter)
2704 | 1299161 | Silicon | Detector
2705 | 1299161 | ±1 | Calibration Uncertainty (%)(k=2)
2706 | 1299161 | ±1 | Power Linearity (%)
2707 | 1299161 | 514 | Calibration Wavelength (nanometers)
2708 | 1299161 | Air | Cooling Method
2709 | 1299161 | 2.5 | Cable Length (meters)
2710 | 1299161 | None | Window
2711 | 1299161 | Yes | Diffuser
2712 | 1299161 | ±4 (325 to 900 nm), ±5 (900 to 1065 nm) | Spectral Compensation Accuracy (%)
2713 | 1299161 | USB | Cable Type
2714 | 1299161 | 140 mW | Maximum Power
2715 | 1299161 | 8.5 µW | Minimum Power
2716 | 1299161 | 325 nm to 1065 nm | Wavelength Range
2717 | 1299161 | 8 mm | Active Area Diameter
2718 | 1299161 | Yes | C24
2719 | 1305568 | 80W, air-cooled <5 min. | Maximum Intermittent Power (watts)(<5 minutes)
2720 | 1305568 | Water, Air (intermittent) | Cooling Type
2721 | 1305568 | RS232 | Cable Type
2722 | 1305568 | 150 | Minimum Wavelength (nanometers)
2723 | 1305568 | 1000 | Maximum Wavelength (nanometers)
2724 | 1305568 | 0.3 | Minimum Power (watts)
2725 | 1305568 | 150 | Maximum Power (watts)
2726 | 1305568 | 150 | Maximum Long-Pulse Energy (joules)
2727 | 1305568 | 0.15 to 1 µm | Wavelength Range (micrometers)
2728 | 1305568 | 300 mW to 150W | Power Range
2729 | 1305568 | 1 to 150 | Long-Pulse Joules Range (joules)
2730 | 1305568 | 1 | Noise Equivalent Power (milliwatts)
2731 | 1305568 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2732 | 1305568 | 600 | Maximum Energy Density (millijoules per square centimeter)
2733 | 1305568 | ±3 | Calibration Uncertainty (%)(k=2)
2734 | 1305568 | ±1 | Power Linearity (%)
2735 | 1305568 | ±1.5 | Spectral Compensation Accuracy (%)
2736 | 1305568 | 514 | Calibration Wavelength (nanometers)
2737 | 1305568 | Water | Cooling Method
2738 | 1305568 | 0.3 | Cable Length (meters)
2739 | 1305568 | None | Window
2740 | 1305568 | None | Diffuser
2741 | 1305568 | 50 mm | Active Area Diameter
2742 | 1305568 | 3 | Response Time, Speed-up On, 0% to 95% (seconds)
2743 | 1305568 | UV | Detector Coating
2744 | 1317151 | 600 | Maximum Energy Density (millijoules per square centimeter)
2745 | 1317151 | RS232 | Cable Type
2746 | 1317151 | 190 | Minimum Wavelength (nanometers)
2747 | 1317151 | 11000 | Maximum Wavelength (nanometers)
2748 | 1317151 | 0.3 | Minimum Power (watts)
2749 | 1317151 | 150 | Maximum Power (watts)
2750 | 1317151 | 150 | Maximum Long-Pulse Energy (joules)
2751 | 1317151 | Air | Cooling Type
2752 | 1317151 | 0.19 to 11 µm | Wavelength Range (micrometers)
2753 | 1317151 | 300 mW to 150W | Power Range
2754 | 1317151 | 300 | Maximum Intermittent Power (watts)(<5 minutes)
2755 | 1317151 | 1 to 150 | Long-Pulse Joules Range (joules)
2756 | 1317151 | 1 | Noise Equivalent Power (milliwatts)
2757 | 1317151 | 6 kW/cm² | Maximum Power Density (kilowatts per square centimeter)
2758 | 1317151 | Broadband | Detector Coating
2759 | 1317151 | ±2 | Calibration Uncertainty (%)(k=2)
2760 | 1317151 | ±1 | Power Linearity (%)
2761 | 1317151 | ±1.5 | Spectral Compensation Accuracy (%)
2762 | 1317151 | 10600 | Calibration Wavelength (nanometers)
2763 | 1317151 | Air | Cooling Method
2764 | 1317151 | 0.3 | Cable Length (meters)
2765 | 1317151 | None | Window
2766 | 1317151 | None | Diffuser
2767 | 1317151 | 50 mm | Active Area Diameter
2768 | 1317151 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2769 | 1360550 | Continuous (CW), pulsed | Capture Modes
2770 | 1360550 | 1 msec to 500 msec, default at 5 msec | Variable Exposure Time
2772 | 1360550 | 190 | Minimum Wavelength (nanometers)
2773 | 1360550 | 355 | Maximum Wavelength (nanometers)
2774 | 1360550 | 0.5 to 6.0 | Recommended Beam Diameters (millimeters)
2775 | 1360550 | 14-bit | Camera Bit Depth
2776 | 1360550 | 1280 x 1024 | Sensor Elements (pixels)
2777 | 1360550 | 20x20 | Effective Pixel Resolution (micrometers)
2778 | 1360550 | n/a | Pixel Size (micrometers)
2779 | 1360550 | 8.3 x 6.6 | Sensor Active Area (millimeters)(H x V)
2780 | 1360550 | 190 to 355 | Spectral Range (nanometers) without LDFP
2781 | 1360550 | 20 | Trigger Delay (microseconds)
2782 | 1360550 | 200 | Maximum Pulse Trigger in Rate (hertz)
2783 | 1360550 | .2 mJ/cm2 at 1064 nm | Damage Threshold without LDFP
2784 | 1360550 | 90 mW/cm2 at 248 nm | CW Saturation with LDFP at 633 nanometers
2785 | 1360550 | .09 mW/cm2 at 248 nm | CW Saturation without LDFP at 633 nanometers
2786 | 1360550 | 5 mJ/cm2 at 248 nm | Pulsed Saturation with LDFP
2787 | 1360550 | .005 mJ/cm2 at 248 nm | Pulsed Saturation without LDFP
2788 | 1360550 | Connector BNC receptacle (trigger cable included) | Trigger
2789 | 1360550 | 190 nm to 355 nm | Wavelength Range
2790 | 1360550 | 500 microns to 6 mm | Recommended Beam Size
2791 | 1360550 | Camera (CW or Pulsed) | Profiler Type
2792 | 1360550 | 8.3 x 6.6 mm (2/3 inch format) | Active Area Size
2793 | 1360550 | LCAM20 | Coupon Spec
2794 | 1402728 | 190 | Minimum Wavelength (nanometers)
2795 | 1402728 | 11000 | Maximum Wavelength (nanometers)
2796 | 1402728 | 10 | Minimum Power (watts)
2797 | 1402728 | 6000 | Maximum Power (watts)
2798 | 1402728 | 50 mm | Active Area Diameter
2799 | 1402728 | Water | Cooling Type
2800 | 1402728 | PM DB-25 | Cable Type
2801 | 1402728 | 0.19 to 11 | Wavelength Range (micrometers)
2802 | 1402728 | 10W to 6000W | Power Range
2803 | 1402728 | 6000 | Maximum Intermittent Power (watts)(<5 minutes)
2804 | 1402728 | 100 | Noise Equivalent Power (milliwatts)
2805 | 1402728 | 2.3 to 14 | Maximum Power Density (kilowatts per square centimeter)
2806 | 1402728 | 600 | Maximum Energy Density (millijoules per square centimeter)
2807 | 1402728 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2808 | 1402728 | BB+ | Detector Coating
2809 | 1402728 | ±3 | Calibration Uncertainty (%)(k=2)
2810 | 1402728 | ±2 | Power Linearity (%)
2811 | 1402728 | ±1.5 | Spectral Compensation Accuracy (%)
2812 | 1402728 | 1070, 10600 | Calibration Wavelength (nanometers)
2813 | 1402728 | Water | Cooling Method
2814 | 1402728 | 2 | Cable Length (meters)
2815 | 1402728 | None | Window
2816 | 1402728 | None | Diffuser
2817 | 1402728 | PM6K20 | Coupon Spec
2818 | 1402729 | 190 | Minimum Wavelength (nanometers)
2819 | 1402729 | 11000 | Maximum Wavelength (nanometers)
2820 | 1402729 | 10 | Minimum Power (watts)
2821 | 1402729 | 6000 | Maximum Power (watts)
2822 | 1402729 | 50 mm | Active Area Diameter
2823 | 1402729 | Water | Cooling Type
2824 | 1402729 | USB | Cable Type
2825 | 1402729 | USB Direct | Meter Compatibility
2826 | 1402729 | 0.19 to 11 | Wavelength Range (micrometers)
2827 | 1402729 | 10W to 6000W | Power Range
2828 | 1402729 | 6000 | Maximum Intermittent Power (watts)(<5 minutes)
2829 | 1402729 | 100 | Noise Equivalent Power (milliwatts)
2830 | 1402729 | 2.3 to 14 | Maximum Power Density (kilowatts per square centimeter)
2831 | 1402729 | 600 | Maximum Energy Density (millijoules per square centimeter)
2832 | 1402729 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2833 | 1402729 | BB+ | Detector Coating
2834 | 1402729 | ±3 | Calibration Uncertainty (%)(k=2)
2835 | 1402729 | ±2 | Power Linearity (%)
2836 | 1402729 | ±1.5 | Spectral Compensation Accuracy (%)
2837 | 1402729 | 1070, 10600 | Calibration Wavelength (nanometers)
2838 | 1402729 | Water | Cooling Method
2839 | 1402729 | 2 | Cable Length (meters)
2840 | 1402729 | None | Window
2841 | 1402729 | None | Diffuser
2842 | 1402729 | PM6K20 | Coupon Spec
2843 | 1402730 | 190 | Minimum Wavelength (nanometers)
2844 | 1402730 | 11000 | Maximum Wavelength (nanometers)
2845 | 1402730 | 10 | Minimum Power (watts)
2846 | 1402730 | 6000 | Maximum Power (watts)
2847 | 1402730 | 50 mm | Active Area Diameter
2848 | 1402730 | Water | Cooling Type
2849 | 1402730 | RS232 | Cable Type
2850 | 1402730 | RS-232 Direct | Meter Compatibility
2851 | 1402730 | 0.19 to 11 | Wavelength Range (micrometers)
2852 | 1402730 | 10W to 6000W | Power Range
2853 | 1402730 | 6000 | Maximum Intermittent Power (watts)(<5 minutes)
2854 | 1402730 | 100 | Noise Equivalent Power (milliwatts)
2855 | 1402730 | 2.3 to 14 | Maximum Power Density (kilowatts per square centimeter)
2856 | 1402730 | 600 | Maximum Energy Density (millijoules per square centimeter)
2857 | 1402730 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2858 | 1402730 | BB+ | Detector Coating
2859 | 1402730 | ±3 | Calibration Uncertainty (%)(k=2)
2860 | 1402730 | ±2 | Power Linearity (%)
2861 | 1402730 | ±1.5 | Spectral Compensation Accuracy (%)
2862 | 1402730 | 1070, 10600 | Calibration Wavelength (nanometers)
2863 | 1402730 | Water | Cooling Method
2864 | 1402730 | 2 | Cable Length (meters)
2865 | 1402730 | None | Window
2866 | 1402730 | None | Diffuser
2867 | 1402730 | PM6K20 | Coupon Spec
2868 | 1409621 | 190 | Minimum Wavelength (nanometers)
2869 | 1409621 | 11000 | Maximum Wavelength (nanometers)
2870 | 1409621 | 5 | Minimum Power (watts)
2871 | 1409621 | 1000 | Maximum Power (watts)
2872 | 1409621 | 50 mm | Active Area Diameter
2873 | 1409621 | Water | Cooling Type
2874 | 1409621 | PM DB-25 | Cable Type
2875 | 1409621 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMate, FieldMaxII-TO, FieldMaxII-TOP | Meter Compatibility
2876 | 1409621 | 0.19 to 11 | Wavelength Range (micrometers)
2877 | 1409621 | 5W to 1000W | Power Range
2878 | 1409621 | 2000 | Maximum Intermittent Power (watts)(<5 minutes)
2879 | 1409621 | 100 | Noise Equivalent Power (milliwatts)
2880 | 1409621 | 10 to 20 | Maximum Power Density (kilowatts per square centimeter)
2881 | 1409621 | 600 | Maximum Energy Density (millijoules per square centimeter)
2882 | 1409621 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2883 | 1409621 | BB+ | Detector Coating
2884 | 1409621 | ±3 | Calibration Uncertainty (%)(k=2)
2885 | 1409621 | ±2 | Power Linearity (%)
2886 | 1409621 | ±1.5 | Spectral Compensation Accuracy (%)
2887 | 1409621 | 1070, 10600 | Calibration Wavelength (nanometers)
2888 | 1409621 | Water | Cooling Method
2889 | 1409621 | 2 | Cable Length (meters)
2890 | 1409621 | None | Window
2891 | 1409621 | None | Diffuser
2892 | 1409622 | 190 | Minimum Wavelength (nanometers)
2893 | 1409622 | 11000 | Maximum Wavelength (nanometers)
2894 | 1409622 | 5 | Minimum Power (watts)
2895 | 1409622 | 1000 | Maximum Power (watts)
2896 | 1409622 | 50 mm | Active Area Diameter
2897 | 1409622 | Water | Cooling Type
2898 | 1409622 | USB | Cable Type
2899 | 1409622 | USB Direct | Meter Compatibility
2900 | 1409622 | 0.19 to 11 | Wavelength Range (micrometers)
2901 | 1409622 | 5W to 1000W | Power Range
2902 | 1409622 | 2000 | Maximum Intermittent Power (watts)(<5 minutes)
2903 | 1409622 | 100 | Noise Equivalent Power (milliwatts)
2904 | 1409622 | 10 to 20 | Maximum Power Density (kilowatts per square centimeter)
2905 | 1409622 | 600 | Maximum Energy Density (millijoules per square centimeter)
2906 | 1409622 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2907 | 1409622 | BB+ | Detector Coating
2908 | 1409622 | ±3 | Calibration Uncertainty (%)(k=2)
2909 | 1409622 | ±2 | Power Linearity (%)
2910 | 1409622 | ±1.5 | Spectral Compensation Accuracy (%)
2911 | 1409622 | 1070, 10600 | Calibration Wavelength (nanometers)
2912 | 1409622 | Water | Cooling Method
2913 | 1409622 | 2 | Cable Length (meters)
2914 | 1409622 | None | Window
2915 | 1409622 | None | Diffuser
2916 | 1409623 | 190 | Minimum Wavelength (nanometers)
2917 | 1409623 | 11000 | Maximum Wavelength (nanometers)
2918 | 1409623 | 5 | Minimum Power (watts)
2919 | 1409623 | 1000 | Maximum Power (watts)
2920 | 1409623 | 50 mm | Active Area Diameter
2921 | 1409623 | Water | Cooling Type
2922 | 1409623 | RS232 | Cable Type
2923 | 1409623 | RS-232 Direct | Meter Compatibility
2924 | 1409623 | 0.19 to 11 | Wavelength Range (micrometers)
2925 | 1409623 | 5W to 1000W | Power Range
2926 | 1409623 | 2000 | Maximum Intermittent Power (watts)(<5 minutes)
2927 | 1409623 | 100 | Noise Equivalent Power (milliwatts)
2928 | 1409623 | 10 to 20 | Maximum Power Density (kilowatts per square centimeter)
2929 | 1409623 | 600 | Maximum Energy Density (millijoules per square centimeter)
2930 | 1409623 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2931 | 1409623 | BB+ | Detector Coating
2932 | 1409623 | ±3 | Calibration Uncertainty (%)(k=2)
2933 | 1409623 | ±2 | Power Linearity (%)
2934 | 1409623 | ±1.5 | Spectral Compensation Accuracy (%)
2935 | 1409623 | 1070, 10600 | Calibration Wavelength (nanometers)
2936 | 1409623 | Water | Cooling Method
2937 | 1409623 | 2 | Cable Length (meters)
2938 | 1409623 | None | Window
2939 | 1409623 | None | Diffuser
2940 | 1409627 | 190 | Minimum Wavelength (nanometers)
2941 | 1409627 | 11000 | Maximum Wavelength (nanometers)
2942 | 1409627 | 5 | Minimum Power (watts)
2943 | 1409627 | 3000 | Maximum Power (watts)
2944 | 1409627 | 50 mm | Active Area Diameter
2945 | 1409627 | Water | Cooling Type
2946 | 1409627 | PM DB-25 | Cable Type
2947 | 1409627 | LabMax-Pro SSIM, LabMax Touch, LabMax Touch Pro, FieldMate, FieldMaxII-TO, FieldMaxII-TOP | Meter Compatibility
2948 | 1409627 | 0.19 to 11 | Wavelength Range (micrometers)
2949 | 1409627 | 5W to 3000W | Power Range
2950 | 1409627 | 3000 | Maximum Intermittent Power (watts)(<5 minutes)
2951 | 1409627 | 100 | Noise Equivalent Power (milliwatts)
2952 | 1409627 | 3.8 to 12 | Maximum Power Density (kilowatts per square centimeter)
2953 | 1409627 | 600 | Maximum Energy Density (millijoules per square centimeter)
2954 | 1409627 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2955 | 1409627 | BB+ | Detector Coating
2956 | 1409627 | ±3 | Calibration Uncertainty (%)(k=2)
2957 | 1409627 | ±2 | Power Linearity (%)
2958 | 1409627 | ±1.5 | Spectral Compensation Accuracy (%)
2959 | 1409627 | 1070, 10600 | Calibration Wavelength (nanometers)
2960 | 1409627 | Water | Cooling Method
2961 | 1409627 | 2 | Cable Length (meters)
2962 | 1409627 | None | Window
2963 | 1409627 | None | Diffuser
2964 | 1409628 | 190 | Minimum Wavelength (nanometers)
2965 | 1409628 | 11000 | Maximum Wavelength (nanometers)
2966 | 1409628 | 5 | Minimum Power (watts)
2967 | 1409628 | 3000 | Maximum Power (watts)
2968 | 1409628 | 50 mm | Active Area Diameter
2969 | 1409628 | Water | Cooling Type
2970 | 1409628 | USB | Cable Type
2971 | 1409628 | USB Direct | Meter Compatibility
2972 | 1409628 | 0.19 to 11 | Wavelength Range (micrometers)
2973 | 1409628 | 5W to 3000W | Power Range
2974 | 1409628 | 3000 | Maximum Intermittent Power (watts)(<5 minutes)
2975 | 1409628 | 100 | Noise Equivalent Power (milliwatts)
2976 | 1409628 | 3.8 to 12 | Maximum Power Density (kilowatts per square centimeter)
2977 | 1409628 | 600 | Maximum Energy Density (millijoules per square centimeter)
2978 | 1409628 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
2979 | 1409628 | BB+ | Detector Coating
2980 | 1409628 | ±3 | Calibration Uncertainty (%)(k=2)
2981 | 1409628 | ±2 | Power Linearity (%)
2982 | 1409628 | ±1.5 | Spectral Compensation Accuracy (%)
2983 | 1409628 | 1070, 10600 | Calibration Wavelength (nanometers)
2984 | 1409628 | Water | Cooling Method
2985 | 1409628 | 2 | Cable Length (meters)
2986 | 1409628 | None | Window
2987 | 1409628 | None | Diffuser
2988 | 1409629 | 190 | Minimum Wavelength (nanometers)
2989 | 1409629 | 11000 | Maximum Wavelength (nanometers)
2990 | 1409629 | 5 | Minimum Power (watts)
2991 | 1409629 | 3000 | Maximum Power (watts)
2992 | 1409629 | 50 mm | Active Area Diameter
2993 | 1409629 | Water | Cooling Type
2994 | 1409629 | RS232 | Cable Type
2995 | 1409629 | RS-232 Direct | Meter Compatibility
2996 | 1409629 | 0.19 to 11 | Wavelength Range (micrometers)
2997 | 1409629 | 5W to 3000W | Power Range
2998 | 1409629 | 3000 | Maximum Intermittent Power (watts)(<5 minutes)
2999 | 1409629 | 100 | Noise Equivalent Power (milliwatts)
3000 | 1409629 | 3.8 to 12 | Maximum Power Density (kilowatts per square centimeter)
3001 | 1409629 | 600 | Maximum Energy Density (millijoules per square centimeter)
3002 | 1409629 | 5 | Response Time, Speed-up On, 0% to 95% (seconds)
3003 | 1409629 | BB+ | Detector Coating
3004 | 1409629 | ±3 | Calibration Uncertainty (%)(k=2)
3005 | 1409629 | ±2 | Power Linearity (%)
3006 | 1409629 | ±1.5 | Spectral Compensation Accuracy (%)
3007 | 1409629 | 1070, 10600 | Calibration Wavelength (nanometers)
3008 | 1409629 | Water | Cooling Method
3009 | 1409629 | 2 | Cable Length (meters)
3010 | 1409629 | None | Window
3011 | 1409629 | None | Diffuser
3012 | 2256258 | Yes | External Trigger Input
3013 | 2256258 | Yes | Battery Operation
3014 | 2256258 | 25000 | Maximum Pyroelectric Rep. Rate (pulses per second)
3015 | 2256258 | ±1.0% | Calibration Uncertainty (%)(k=2)
3016 | 2256258 | Yes | Thermopile sensor compatibility
3017 | 2256258 | Yes | Optical sensor compatibility
3018 | 2256258 | Yes | Pyroelectric sensor compatibility
3019 | 2256258 | Yes | PowerMax-Pro sensor compatibility
3020 | 2256258 | Sensor dependent (reference sensor specifications) | Measurement Range
3021 | 2256258 | Meter accuracy + sensor accuracy | Accuracy - System
3022 | 2256258 | ±0.04% (25% to full scale), ±0.26% (5% to 25% of full scale), ±1% (min to 5% of full scale) | Accuracy - Analog Output
3023 | 2256258 | 10 Hz | Power Sampling Rate
3024 | 2256258 | Yes | Ethernet Interface
3025 | 2256258 | Yes | USB Interface
3026 | 2256258 | Yes | RS232 Interface
3027 | 2256258 | Yes | Analog output
3028 | 2256258 | 7€ TFT LCD with LED backlight, 1024 x 600 pixels. | Display
3029 | 2256258 | External and internal trigger (selectable) | Trigger
3030 | 2256258 | Min., max., mean, range, standard deviation. Trend charting, tuning, positional display, and analysis of beam stability | Statistical Analysis
3031 | 2256258 | 238 x 162 x 51 mm (9.4 x 6.4 x 2.0 in.) | Size (H x W x D)
3032 | 2256258 | 1.19 kg (2.6 lbs.) | Weight
3033 | 2256258 | 12VDC - center positive; (90 to 260 VAC, 50/60 Hz AC power adapter included) | Instrument Power
3034 | 2256258 | CE, RoHS, WEEE, ISO 17025 | Compliance
3035 | 2256258 | Ethernet, USB, RS-232, PC application | Computer Interface
3036 | 2256258 | 0 to 1, 2, or 4 VDC (selectable) | Analog Output Text
3037 | 2256258 | Yes | C24
3038 | 2256259 | Yes | External Trigger Input
3039 | 2256259 | Yes | Battery Operation
3040 | 2256259 | 25000 | Maximum Pyroelectric Rep. Rate (pulses per second)
3041 | 2256259 | ±1.0% | Calibration Uncertainty (%)(k=2)
3042 | 2256259 | Yes | Thermopile sensor compatibility
3043 | 2256259 | Yes | Optical sensor compatibility
3044 | 2256259 | Yes | Pyroelectric sensor compatibility
3045 | 2256259 | Yes | PowerMax-Pro sensor compatibility
3046 | 2256259 | Sensor dependent (reference sensor specifications) | Measurement Range
3047 | 2256259 | Meter accuracy + sensor accuracy | Accuracy - System
3048 | 2256259 | ±0.04% (25% to full scale), ±0.26% (5% to 25% of full scale), ±1% (min to 5% of full scale) | Accuracy - Analog Output
3049 | 2256259 | 10 Hz thermopile and optical, 10 Hz to 10 MHz (user adjustable) with PowerMax-Pro | Power Sampling Rate
3050 | 2256259 | Yes | Ethernet Interface
3051 | 2256259 | Yes | USB Interface
3052 | 2256259 | Yes | RS232 Interface
3053 | 2256259 | Yes | Analog output
3054 | 2256259 | 7€ TFT LCD with LED backlight, 1024 x 600 pixels. | Display
3055 | 2256259 | External and internal trigger (selectable) | Trigger
3056 | 2256259 | Min., max., mean, range, standard deviation. Trend charting, tuning. | Statistical Analysis
3057 | 2256259 | 238 x 162 x 51 mm (9.4 x 6.4 x 2.0 in.) | Size (H x W x D)
3058 | 2256259 | 1.19 kg (2.6 lbs.) | Weight
3059 | 2256259 | 12VDC - center positive; (90 to 260 VAC, 50/60 Hz AC power adapter included) | Instrument Power
3060 | 2256259 | CE, RoHS, WEEE, ISO 17025 | Compliance
3061 | 2256259 | Ethernet, USB, RS-232, PC application | Computer Interface
3062 | 2256259 | 0 to 1, 2, or 4 VDC (selectable) | Analog Output Text
3063 | 2256259 | Yes | C24